# Purpose

To prepare DNA extracted samples for conventional PCR by making a master mix, aliquoting it, adding DNA samples (template), and programming the thermocycler to perform PCR reactions.

# Scope

Making 25 or 50 microliter PCR reactions with DNA extractions. PCR reactions prepared with this SOP will be ready for Sanger sequencing.

# Responsibilities

Those under the supervision of trained personnel.

# Definitions

- Aliquot: Distributing the same amount of a master mix into smaller portions.

- Denaturation: Separating strands of DNA template.

- Annealing: Adding primers to separate strands of DNA.

- Extension: Adding nucleotides with assistance of primers and polymerase to elongate two strands of DNA.

- PCR (Polymerase Chain Reaction): A technique used to amplify a segment of DNA by generating thousands to millions of copies.

- Master Mix: A pre-prepared solution containing most of the reagents needed for PCR, including DNA polymerase, dNTPs, MgCl2, and buffer.

- DNA Template: The DNA sample that contains the target sequence to be amplified in PCR.

- Thermocycler: An instrument used to perform PCR by cycling through different temperature steps.

- Minicentrifuge: A small centrifuge used to spin down small volumes of liquid in PCR tubes.

- GoTaq G2 Colorless Master Mix 2X: A PCR master mix from Promega that contains Taq DNA polymerase, dNTPs, MgCl2, and reaction buffers.

- Upstream Primer: A short single-stranded DNA sequence that binds to the template strand and initiates DNA synthesis during PCR.

- Downstream Primer: A short single-stranded DNA sequence that binds to the complementary template strand and initiates DNA synthesis during PCR.

- Nuclease-free Water: Water that has been treated to remove nucleases, used to prevent degradation of DNA and RNA in experiments.

- ExoSAP-IT™ Express Reagent: A reagent used to clean up PCR products by removing excess primers and dNTPs before sequencing.

- Chill Rack: A rack used to keep PCR tubes cold during sample preparation to prevent degradation.

# References

Promega Usage Information # 9PIM713

# Materials

(Any materials here)

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| PCR tubes or PCR plates | 1 | Bag/Box |
| 1.5-2.0 microcentrifuge tubes | 1 | Bag |
| Micropipettes (p2, p20, p200, p1000) | 1 | Set |
| Pipette tips (for each micropipette type) | 4 | Boxes |
| Thermocycler (e.g. ThermoFisher VeritiPro) | 1 | N/A |
| Minicentrifuge | 1 | N/A |
| GoTaq G2 Colorless Master Mix, 2X | 1 | Tube |
| Upstream primer, 10µm | 1 | Tube |
| Downstream primer, 10µm | 1 | Tube |
| Nuclease-free water | 1 | Tube |
| DNA template | N/A | N/A |
| Chill rack | 1 | N/A |
| ExoSAP-IT™ Express reagent | 1 | Tube |

# Safety

The following PPE is required when working: Nitrile/latex gloves, lab coat, safety glasses, and closed-toe shoes. To avoid contamination: Disinfect workspace with 70% ethanol (wearing PPE) before and after working; replace gloves or apply 70% ethanol on them occasionally.

# Procedure

If preparing 50µL PCR reactions from DNA extracts derived from InstaGene matrix (Chelex)

- Determine number of samples to prepare, including positive control and negative control (n). Then account for 10-20% extra (e.g. 0.2 * n).

- In clean microcentrifuge tube labelled ‘mastermix’ or ‘MM’, add the following with an adjustable micropipette (mix by pipetting and changing tips in between):

  - 25 µL * (n + 0.2n) of GoTaq Colorless Master Mix (2x).

  - 1 µL * (n + 0.2n) of Upstream primer (10µm).

  - 1 µL * (n + 0.2n) of Downstream primer (10µm).

  - 3 µL * (n + 0.2n) of nuclease-free water.

- Aliquot 30µL of master mix into clean set of PCR strips, PCR tubes, or a PCR plate (keep PCR reactions on chill rack while being prepared).

- In same set of PCR strips, PCR tubes, or a PCR plate, add 20µL of each DNA extract obtained from Instagene Matrix (Chelex) based method and mix by pipetting (change tips in between; see relevant SOPs for obtaining 20µL of bead-free supernatant).

- Seal PCR reactions and spin down with minicentrifuge. Keep on chill rack until thermocycler is programmed and ready.

If preparing 25µL PCR reactions from DNA extracts

- Determine number of samples to prepare, including positive control, negative control and 10-20% extra.

- In clean microcentrifuge tube labelled ‘mastermix’ or ‘MM’, add the following with an adjustable micropipette (mix by pipetting and changing tips in between):

  - 12.5 µL * (n + 0.2n) of GoTaq Colorless Master Mix (2x).

  - 1 µL * (n + 0.2n) of Upstream primer (10µm).

  - 1 µL * (n + 0.2n) of Downstream primer (10µm).

  - 5.5 µL * (n + 0.2n) of nuclease-free water.

- Aliquot 20µL of master mix into clean set of PCR strips, PCR tubes, or a PCR plate (PCR reactions on chill rack).

- In same set of PCR strips, PCR tubes, or a PCR plate, add 5µL of each DNA extraction sample (mix by pipetting and change tips in between).

- Seal PCR reactions and spin down with minicentrifuge. Keep on chill rack until thermocycler is programmed and ready.

Operating VeritiPro thermocycler

- Power thermocycler on with switch on back right of instrument.

- Click ‘Sign in’ and use buttons on panel to enter the correct profile and password provided by supervisor. Click ‘Enter’ when done.

- Next, click on ‘Set up run’ panel button followed by ‘Open method’. Click on directory in panel where pre-loaded program exists (consult with supervisor for guidance on where program is or how to create or modify from another).

- Once in the correct directory, select the appropriate ‘Method Name’.

- Confirm with supervisor that settings are correct (cycle temperature, timing, and repetitions; cover, and volume settings. If needed, click on values needing to be changed, use panel values to make adjustments followed by clicking ‘Enter’. (do not save program unless authorization is allowed).

- When done checking and/or adjusting program’s values, click ‘Next’. Ignore prompts to save program unless authorized by instructor.

- Lift lid on top of thermocycler, and place PCR reactions in with caps on tight.

- Close the lid on top of the thermocycler and click the ‘Start run’ button on panel.

Test gel of PCR products prior to sequencing (optional)

See instructions of SOP VG010 for preparing and running agarose gel with PCR samples.

PCR cleanup for Sanger sequencing

- In clean PCR tubes, strips, or PCR plate (sitting on chillrack), transfer 5 µL of PCR product followed by 2 µL of ExoSAP-IT™ Express reagent (changing tips in between and mixing by pipetting).

- Following directions from Operating VeritiPro thermocycler, navigate to ExoSAP-IT™ Express program in VeritiPro thermocycler. Load cleanup reactions into thermocycler and run.

- Proceed to setting up Sanger sequencing reactions or store PCR products in freezer with clear labels of contents and date on plate (or on box for stored PCR tubes).

# Figures and Diagrams

N/A

# Notes/Observations

# Attachments

Appendix – List of recommended Primer-Probe Assays and DNA Controls

VG007 Attachment - Template and Example PCR Master Mix Reaction Sheet