# Purpose

This SOP describes preparation of qPCR reactions, including master mix assembly, template dispensing, and QuantStudio programming for microbial load analysis using four assays: 16S (bacterial load), ITS (fungal load), a taxon-specific assay, and a host housekeeping gene qPCR.

# Scope

This procedure applies to preparation of 20 µL TaqMan qPCR reactions run on a 96‑well optical block using a standardized plate layout with triplicate biological samples.

# Responsibilities

This procedure is to be performed by trained personnel or trainees under supervision.

# Definitions

- Aliquot: Distributing the same amount of a master mix into smaller portions.

- Denaturation: Separating strands of DNA template.

- Annealing: Adding primers to separate strands of DNA.

- Extension: Adding nucleotides with assistance of primers and polymerase to elongate two strands of DNA.

- qPCR (Quantitative Polymerase Chain Reaction): A laboratory technique used to amplify and simultaneously quantify a targeted DNA molecule.

- TaqMan Assay: A specific type of qPCR assay that uses a fluorescent reporter dye to quantify DNA amplification.

- Master Mix: A pre-prepared solution containing most of the reagents needed for PCR, including DNA polymerase, dNTPs, MgCl2, and buffer.

- DNA Template: The DNA sample that contains the target sequence to be amplified in PCR.

- Thermocycler (QuantStudio): An instrument used to perform qPCR by cycling through different temperature steps and detecting fluorescence.

- Minicentrifuge: A small centrifuge used to spin down small volumes of liquid in PCR tubes.

- Optical PCR Tubes/Plates: Tubes or plates designed for use in qPCR, allowing for the detection of fluorescence during the reaction.

- TaqMan Fast Advanced Master Mix 2x: A PCR master mix from Applied Biosystems designed for fast and accurate qPCR.

- Primer-Probe Assay (20x): A mixture of primers and probe used in TaqMan qPCR assays.

- Nuclease-Free Water: Water that has been treated to remove nucleases, used to prevent degradation of DNA and RNA in experiments.

- Human Genomic Control DNA: A DNA sample with a known concentration used as a positive control in qPCR.

- 20X RNase P Primer-Probe (FAM™ Dye) Mix: A primer-probe mix used as a control in qPCR assays.

# References

TaqMan Fast Advanced Master Mix User Guide (Pub No. Man0025706, Rev. A.0)

# Materials

(Any materials here)

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Optical PCR tubes or PCR plate with optical film | N/A | N/A |
| 2.0 mL microcentrifuge tubes | N/A | N/A |
| Micropipettes (p2, p20, p200, p1000) | 1 | Set |
| Pipette tips (for each micropipette type) | 4 | Boxes |
| Thermocycler (e.g. QuantStudio 3) | 1 | N/A |
| Minicentrifuge | 1 | N/A |
| TaqMan Fast Advanced, 2x | 1 | Tube |
| Primer-Probe Assay (20x) [See Appendix] | 1-4 | Tubes |
| DNA Standards [See Appendix] | 1 | Set of 7 |
| Nuclease-free water | N/A | N/A |
| DNA templates / extractions | TBD | N/A |
| Chill rack / ice bucket | 1 | rack/bucket |
| Human Genomic Control DNA 10 ng/µL | N/A | Tube |

# Safety

Nitrile/latex gloves, lab coat, safety glasses, and closed-toe, non-slip shoes are required. Tubes, tips, and plates not in contact with biohazards go in regular waste. Disinfect workspace with 70 percent ethanol before and after working, and replace gloves as needed.

# Procedure

Setup

- Thaw reagents and samples.

- Flick tubes to mix and spin down 1-2 seconds.

- Place and work with temperature sensitive items (e.g. master mix and DNA templates) on chill rack or ice bucket.

- Prepare 20x primer-probe Taqman assay if not already prepared.

  - Reconstitute oligos with nuclease-free water so the selected forward primer, reverse primer, and probe solutions are each 100μM in concentration.

  - Combine 90μL forward primer, 90μL reverse primer, and 25μL probe stocks.

  - Add 295μL nuclease-free water to finalize 20x working stock of Taqman primer-probe assay, mix, spin, and store dark on ice before use (store at -20∘C when done).

- If comparing amounts of amplified DNA targets across unknown samples, ensure that samples are equal in their overall DNA concentration (20 to 100 ng/uL according to instructions of supervisor) and are of good quality.

  - See SOPs VG001 and VG005 for determining DNA sample concentration and quality.

  - Normalize each DNA sample as needed with molecular-grade water (e.g. adding 100 µL of molecular-grade water to a 100 µL of 100 ng/µL DNA sample will normalize the DNA sample to a desired concentration of 50 ng/µL).

  - If microbial amplicons are being normalized analytically (post qPCR experiments) against amplicons of a host housekeeping gene (see qPCR analysis spreadsheet), no dilution/concentration steps are needed for DNA samples.

  - However, DNA samples should be adjusted to be less than 100 ng/µL to minimize inhibitors within DNA sample, and if failure to amplify occurs, samples may need to be further diluted.

Preparing 20µL qPCR reactions

Each reaction (20 µL) requires:
 	• 10 µL TaqMan Fast Advanced Master Mix (2×)
	 • 1 µL 20× Primer/Probe mix
	 • 7 µL nuclease‑free water

Prepare master mix for 96 total reactions to cover all wells specified in the plate layout 	below while planning for >10% excess volume.

- Add the following in a sterile 2mL microcentrifuge tube, mixing by pipetting and switching tips with each new reagent added to master mix:

  - 960 µL TaqMan Fast Advanced Master Mix (2×)

  - 96 µL 20× Primer/Probe mix

  - 672 µL nuclease‑free water

- Aliquot 18µL of master mix into each well assigned in the plate layout below.

Plate Layout

The plate layout below applies identically to all four primer-probe assays. Use “qPCR experiment setup record” attachment to enter sample and reagent names.

|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| A | Std1 | Std2 | Std3 | Std4 | Std5 | Std6 | Std7 | NTC | NTC | EB1 | EB2 |  |
| B | Std1 | Std2 | Std3 | Std4 | Std5 | Std6 | Std7 | NTC |  |  |  |  |
| C | Std1 | Std2 | Std3 | Std4 | Std5 | Std6 | Std7 |  |  |  |  |  |
| D | S1 | S1 | S1 | S2 | S2 | S2 | S3 | S3 | S3 | S4 | S4 | S4 |
| E | S5 | S5 | S5 | S6 | S6 | S6 | S7 | S7 | S7 | S8 | S8 | S8 |
| F | S9 | S9 | S9 | S10 | S10 | S10 | S11 | S11 | S11 | S12 | S12 | S12 |
| G | S13 | S13 | S13 | S14 | S14 | S14 | S15 | S15 | S15 | S16 | S16 | S16 |
| H | S17 | S17 | S17 | S18 | S18 | S18 | S19 | S19 | S19 | S20 | S20 | S20 |

Dispensing DNA Template

Add 2 µL of the appropriate template into each well:

- Std1–Std7 into wells A1–C7

- NTCs into A8, A9, B8

- EB1 and EB2 into A10 and A11

- Biological samples S1–S20 into triplicate groups in rows D–H

- Seal plate with optical film and spin down 2–3 seconds.

Edit an existing QuantStudio program

- Turn on laptop connected to QuantStudio.

- Login with profile and password provided by supervisor.

- Click on desktop shortcut for ‘QuantStudio Design & Analysis’ program.

- Click on ‘Open Existing Experiment’. Navigate to and click ‘open’ on an existing template (.edt) file saved on the laptop (ask supervisor where to locate directory containing .edt files).

- Once template file opens and displays the ‘Properties’ tab, adjust the following fields and dropdowns for quantification of normalized sets DNA samples using a TaqMan probe assay.

  - Instrument type: QuantStudio 3 System.

  - Block type: 96-Well 0.2-mL Block.

  - Experiment type: Standard Curve.

  - Chemistry: TaqMan Reagents.

  - Run mode: Standard.

- Next go to ‘Method’ tab to and click on the fields/buttons listed below to inspect and/or adjust values of the qPCR settings:

  - Volume: 20 µL.

  - Cover: 105 °C.

  - Hold Stage (Step 1): 50.0 °C for 2 minutes. 1.6 °C/s ramp rate.

  - Hold Stage (Step 2): 95.0 °C for 2 minutes. 1.6 °C/s ramp rate.

  - PCR Stage set for 40x cycles.

  - PCR Stage (Step 1): 95.0 °C for 1 second. 1.6 °C/s ramp rate.

  - PCR Stage (Step 2): 60.0 °C for 20 seconds. 1.6 °C/s ramp rate.

  - Click on camera button of PCR Stage (Step 2). Make sure it is off for other steps.

- Next go to ‘Plate’ tab to specify/modify sample-reagent details by clicking on a plate position and adjusting the fields/buttons listed below:

  - Target: Name of TaqMan probe chemistry (including its reporter and quencher) that you are planning to use for a given sample.

Hint: To setup a new target and its details, go to ‘Advanced Setup’, click the + symbol, and choose from the dropdown options that match your target.

  - Sample: Name of the reaction you have setup for a given plate postion.

  - Task: Specifies whether to treat a given reaction as an unknown sample, a control sample, or a standard.

- Once all settings are specified, go to ‘Run’ tab, click on the ‘Save’ and ‘Save as’ button, and choose name and directory/USB where to save the template (.edt) file.

Operating QuantStudio 3 thermocycler

- Turn on QuantStudio using switch on back right of thermocycler.

- Click on the ‘Sign in’ button on panel.

- Choose the user name associated with your profile, and enter the PIN provided by the supervisor using buttons on the panel. Click ‘Enter’ when done.

- Click the up-arrow on the top-right of the home screen to open the tray.

- Insert the PCR plate with your prepared reactions or insert PCR strips or tubes loaded into the PCR plate adaptor (white base and top with tubes/strips sandwiched between).

- Make sure that optical adhesive film is completely covering wells of plate with all surfaces smooth and clear (use a plastic edge to run over film to avoid finger prints). If using tubes or strips, make sure optical caps are firmly sealed.

- Click the up-arrow on panel again to close the tray with reactions loaded.

- Now click ‘Set up run’ button on middle of home screen.

- Make sure computer is hooked up to thermocycler or USB containing the appropriate .edt file is connected to USB port of thermocycler.

- Click ‘Load Experiment’ and follow the prompts on screen to specify where the desired template (.edt) file is.

- Once the file is clicked open, review the settings on the Method and Plate buttons on the screen to confirm that you have loaded the correct .edt file corresponding to your plate setup, sample, and target settings.

- Click ‘Start Run’ to run qPCR experiment.

Exporting Ct data for analysis

- When the thermocycler has finished, you will be prompted on the panel where to save the .eds file.

- Transfer the .eds file to the computer with the ‘QuantStudio Design & Analysis’ program.

- Click the ‘Open Existing Experiment’ button on ‘QuantStudio Design & Analysis’ program and choose the .eds file transferred from the QuantStudio.

- Click the ‘Export’ button, type in the ‘File Name’ field what you want to call the exported table.

- Choose .xls in the ‘File Type’ dropdown and save/export the file to a directory of your choice. Copy and/or transfer the .xls file to a computer capable of running Excel.

- Copy and paste the Ct values corresponding to each reaction into the appropriate fields of the qPCR analysis spreadsheet (see attachments). Summary statistics and bar graphs will automatically generate results from the raw data transferred onto the spreadsheet.

# Figures and Diagrams

N/A

# Notes/Observations

N/A

# Attachments

- Appendix – List of recommended primer, probes, and DNA Controls

- VG009 Attachment - qPCR experiment setup record

- VG009 Attachment - qPCR Analysis Spreadsheet