# Purpose

To perform Sanger cycle sequencing on PCR-amplified products.

# Scope

Perform cycle sequencing on PCR samples after enzymatic cleanup.

# Responsibilities

Those under the supervision of trained personnel.

# Definitions

- Sample number = n

- Aliquot: Distributing the same amount of a master mix into smaller portions.

- Sequencing Batch Document: A document used to track the number of samples and their respective combinations of PCR samples and primer-containing master mixes in a sequencing experiment.

- Master Mix (MM): A prepared solution containing all the reagents necessary for a sequencing reaction, except for the template DNA.

- BigDye Reaction Mix: A commercial mixture used in sequencing reactions that includes DNA polymerase, dNTPs (dATP, dCTP, dGTP, dTTP), and dye-labeled ddNTPs.

- Nuclease-Free Water: Highly purified water free from DNases and RNases used in molecular biology experiments to prevent degradation of nucleic acids.

- PCR Strips/Tubes/Plates: Containers used to hold samples during PCR and sequencing reactions. They are designed to fit into thermal cyclers and sequencing instruments.

- Thermocycler: An instrument used to amplify DNA and perform other thermal cycling techniques, essential for PCR and sequencing reactions.

- SeqStudio: A sequencing instrument used for analyzing DNA samples through capillary electrophoresis.

- Cathode Buffer Cartridge: A component of the SeqStudio that contains the buffer solution needed for the electrophoresis process.

- Plate Manager Program: Software used to manage sample information and setup for sequencing on the SeqStudio.

- SAM Solution: A solution used to prepare sequencing reactions for cleanup by binding to excess dyes and primers.

- BigDye XTerminator Bead Solution: Beads used in the cleanup process to remove unincorporated dye terminators from sequencing reactions.

- .AB1 Files: Files that contain data from automated DNA sequencing instruments. These files include the raw data output from the sequencer.

# References

- BigDye Terminator v3.1 Cycle Sequencing Kit User Guide.

- SeqStudio Genetic Analyzer Instrument and Software User Guide

# Materials

(Any materials here)

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| BigDye Terminator 3.1 Ready Reaction Mix (2x) | 1 | tube |
| Upstream primer (10µM) | 1 | tube |
| Downstream primer (10µM) | 1 | tube |
| Deionized, nuclease-free water | 1 | tube |
| DNA templates | TBD | 20-100ng/µL samples |
| Positive control (pGEM) | 1 | tube |
| M13 primer for positive control | 1 | tube |
| Microcentrifuge Tubes | 1 | bag |
| Adjustable Micropipettes and Tips | 1 | set |
| PCR Strips, Tubes, or Plates | 1 | bag/box |
| Thermocycler (e.g., VeritiPro) | 1 | N/A |
| SeqStudio Sequencer | 1 | N/A |
| Cathode Buffer | 1 | container |
| Flow Cell Cartridge | 1 | N/A |
| SAM Solution | 1 | tube |
| BigDye XTerminator Bead Solution | 1 | tube |
| Plate Shaker or Digital Vortex | 1 | N/A |
| Centrifuge (e.g. Minicentrifuge) | 1 | N/A |
| Septa Strips | 1 | bag/box |
| Laptop with Plate Manager Software | 1 | N/A |
| External Storage or Network Drive | 1 | N/A |

# Safety

The following PPE is required when working: Nitrile/latex gloves, lab coat, safety glasses, and closed-toe shoes.

# Procedure

Setting up Sanger cycle sequencing reactions

- Using Sequencing Batch Document, determine number of samples to prepare, including upstream reactions, downstream reactions, control and 10-20% extra and note the desired combination of PCR samples and primer-containing master mixes used for each reaction tube/well.

- For each sequencing master mix containing a unique primer, add the following in a clean microcentrifuge tube labelled ‘mastermix’ or ‘MM’ (use adjustable micropipette transfer and mix while changing tips in between):

  - 4 µL * (n + 0.2n) of BigDye Reaction Mix (2x)

  - 1 µL * (n + 0.2n) of Upstream or Downstream primer (10µm)

  - 4 µL * (n + 0.2n) of nuclease-free water

  - Aliquot 9µL of master mix into clean set of PCR strips, PCR tubes, or a PCR plate (PCR reactions on chill rack)

- Repeat step 2 if preparing another sequencing master mix with different primer.

- In same set of PCR strips, PCR tubes, or a PCR plate, add 1µL of each unique PCR sample to its designated aliquot of sequencing reaction master mix (mix by pipetting and change tips in between).

- Prepare at least one positive control in a PCR tube by transferring the following (and changing tips in between): 4 µL BigDye Reaction Mix, 1 µL M13 primer, 4 µL nuclease-free water, and 1 µL of pGEM template.

- Seal sequencing reactions and spin down 1-2 seconds with minicentrifuge (~1,000 rpm). Keep on chill rack until thermocycler is programmed and ready.

Operating VeritiPro thermocycler

Power on with switch on back of thermocycler.

On panel, click ‘Open template’ and then ‘Big Dye Direct Sequencing 3.0’

Lift open lid of thermocycler, place sealed sequencing reactions in, and close lid.

Click ‘Run’.

Preparing SeqStudio sequencer

- Turn on SeqStudio with switch on back of instrument.

- Login with profile and password provided by supervisor.

- To check or replace cathode buffer cartridge (if buffer below fill line):

  - Click up arrow button on panel, lift door of instrument, and click ‘Eject Plate’.

  - When tray has moved up front, push gray button to open lid and replace or inspect cathode buffer cartridge with septa inserted (buffer should be above fill line).

- To install cartridge:

  - Click up arrow button on panel and open door of instrument.

  - Remove optical protector and capillary protector on cartridge (place in fridge or discard if used two or more times).

  - Insert cartridge with arrow on top pointing in appropriate insert of instrument (make sure it clicks).

  - Click ‘Close’ on panel and close door.

- To setup or edit SeqStudio program:

  - Turn on laptop connected to SeqStudio.

  - Login with profile and password provided by supervisor.

  - Click on desktop shortcut for Plate Manager program.

  - Click on ‘Create a plate file’, then ‘Next’ and Next’.

  - Click on each injection group based on how you want to setup plate.

  - For each reaction, enter sample name, choose dye set from dropdown (Z_BigDye Terminator v3.1), and select run module (Medium SeqBDX for 300-900 base pair fragments).

  - Click ‘Next’ followed by ‘Save’. Designate where to save file.

  - Go to panel of SeqStudio, and choose ‘Import’ to select the plate setup file from the laptop.

Sanger sequencing cleanup

- When thermocycler program has ended, take sequencing reactions out.

- In clean microcentrifuge tube, add 45 µL * (n + 0.2n) of SAM solution.

- In same tube, add 10 µL * (n + 0.2n) of pre-vortexed BigDye XTerminator bead solution.

- Aliquot 55 µL of SAM-BigDye XTerminator bead solution mix to each prepared sequencing reaction (mix by pipetting 3-4 times in between transfers).

- With seal on, shake samples at 1,800 rpm for 20 minutes using plate shaker or digital vortex with adaptor for PCR tubes/strips.

- Centrifuge samples at 1,000 rpm for 2 minutes using approved and properly counterbalanced centrifuge.

- Gently replace lids of samples in PCR tubes with septa strips for each column of samples. (Note: If using PCR tubes or strips, place samples in base of white plate adapter first, followed by addition of top half of white plate adaptor and septa strips for each column of 8 samples).

Operating SeqStudio sequencer

- Click up arrow button on panel, lift door of instrument, and click ‘Eject Plate’

- When tray has moved up front, push gray button to open lid and place tubes in their appropriate insert of the tray.

- Close lid (listen for click).

- Click ‘Retract Plate’ on the panel and close the SeqStudio door.

- After confirming that the placement of samples aligns with the program run settings (e.g. specifying BigDye V3 reagents, MedseqBDX settings, and correct wells utilized), click ‘Start Run’.

Saving and exporting files

- When prompted click ‘Export’ to save report to connected laptop.

- In designated directory where the program file was saved, find the report and .AB1 files for your sequencing experiment.

- From this directory transfer .AB1 chromatogram files to your program or analytical pipeline of preference (e.g. see attachment Analyze and export DNA sequence files on CyVerse).

# Figures and Diagrams

N/A

# Notes/Observations

# Attachments

See Sanger Sequencing Batch Document

Analyze and export DNA sequence files on DNA Subway