| Document No.: BioWork-1030 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 1 of 1 |
| --- | --- | --- | --- |
| Following Written Instructions |  |  |  |

#### Purpose

Practice following written instructions like those found in Standard Operating Procedures (SOPs).

#### Procedure

  - Read all instructions first

  - Write all of your team's initials at the top right hand corner of this sheet

  - Write your first name on your sheet of paper

  - Write the total of 3 + 16 + 32 + 64 here:

  - Underline instruction 2.1 above

  - Check the time by your watch with that of one of your neighbor's

  - Write down the difference in time between the two watches at the foot of this page

  - Draw three circles in the left hand margin

  - Put a tick in each of the circles mentioned in 2.8

  - Sign your signature at the foot of the page

  - On the back of the page, divide 50 by 12.5

  - When you get to this point in the test, stand up, then sit down and continue with the next item

  - If you have carefully followed all these instructions, call out ‘I have’

  - On the reverse of this page, draw quickly what you think an upright bicycle looks like from overhead

  - Check your answer to item 2.4, multiply it by 5 and write the result in the left hand margin opposite this item

  - Write the 5th, 10th, 9th and 20th letters of the alphabet here:

  - Punch three holes with your pen here: o o o

  - If you think you are the first person to get this far, call out ‘I’m in the lead’

  - Underline all the even digits on the left hand side of the page

  - Draw triangles round the holes you punched in Item 2.17

  - Now you’ve finished reading all the instructions, obey only Steps 2.2, 2.10 & 2.22

  - Stand up and say, “We’re the greatest team in the World!”

- Attachments

  - N/A

#### Purpose

This procedure describes the current Good Manufacturing Practice (CGMP) documentation requirements.

#### Scope

This procedure applies to all individuals participating in the BioWork course when recording CGMP-related activities on official and controlled documents.

Note: This is the note section.

#### Responsibilities

It is the employee’s responsibility to make sure that the procedures outlined in this document are followed.

The department supervisor will be responsible for making sure employees are up to date on training and being compliant with the Good Documentation Practices (GDP).

#### Definitions

CGMP:	current Good Manufacturing Practice

CGMP documents:	Lab protocols, lab notebooks, exams and quizzes, and other documents as noted will be treated as a CGMP document.

Copy: Reproduction of an original document.

GDP: Good Documentation Practice

Performer:	The employee who executes and/or documents the execution of a process or step(s) in a process.

Real-Time:	At the time of performance.

Reviewer:	One who examines and/or evaluates the contents of documents to ensure accuracy, consistency, and adherence to procedures after performance.

Verifier:	The employee who has personally observed and therefore can attest to the execution of a process or step(s) in a process.

#### References

- 21 CFR Part 211 – Subpart J – Records and Reports https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfcfr/CFRsearch.cfm?CFRPart=211

#### Materials

N/A

#### Safety

N/A

#### Procedure

  - General Documentation Practices

    - Entries made to CGMP documents (e.g., lab protocols, log books, forms, tests, etc.) must be clear, legible, and recorded in indelible blue ink.

    - Data recorded on CGMP documents must be recorded in real-time or as soon as safely possible following completion of the task.

    - NOTE: Backdating is not allowed under any circumstance nor is signing before you have completed a task.

    - All performers must sign/initial and date the document in the “Performed By/Date” space. If verification is required, the student verifying the task must sign/initial and date the “Verified By/Date” space.

    - NOTE: Never sign a document for another person or log in as another person.

    - Additional entries or comments on documents must be initialed and dated at the time of entry by the individual recording the comment. If verification is necessary in the performance of a referenced task, a second student must also initial and date the comment.

    - All blank sections or data entry lines on documents must be addressed. If no data is applicable, draw a single line across the blank area(s), enter N/A, initial and date.

    - NOTE: If a “N/A” checkbox is provided, the student may check the box, rather than initialing and dating the entry.

    - When recording or reporting numerical data, always ensure that a “0” is used as a placeholder to the left of the decimal point if no other whole number is applicable.

    - The use of ditto marks or arrows pointing downward to indicate repeating data entries is prohibited.

    - Highlighters and correction fluid may not be used on CGMP treated documents such as lab protocols, lab notebooks, and exams.

  - Original Documentation

    - Original CGMP documents must never be destroyed or discarded.

    - If necessary, e.g., in case of a spill, an original document may be copied or transcribed onto a new CGMP document upon supervisor approval. To identify and void the original document, a single diagonal line must be drawn through each page. Annotate on the first page of the original document that “Data has been transcribed due to”, then proceed with an explanation, initial and date the entry. The “copied” or “transcribed” document must be identified as such with reference made to the original document. Attach the original document to the new “copy”.

  - Attachments

    - To ensure document traceability, all attachment pages to original document(s) require cross-reference to original document(s), pagination such as “Attachment 1 of 3” and so on for all attachments, information such as step number as required, and initial and date.

  - Recoding of Dates

    - To ensure consistency in documentation practices, dates must be recorded as DD MMM YY, where D = day, M = month, and Y = year.

    - Example: 06 Apr 17

    - When copying a date from a source (such as calibration due dates) the date must be transcribed as written.

  - Recording of Time

    - Recording of time must be obtained from a common source (e.g., computer, class clock, etc.). Time displayed on sources such as wristwatches must not be used.

    - Twenty-four-hour format or military time is required on all CGMP documentation when prompted.

| Document No.: BioWork-1060 | Revision No.: 01 | Effective Date: 23 Aug 2019 | Page 4 of 5 |
| --- | --- | --- | --- |
| CGMP Documentation Practices |  |  |  |

    - When recording elapsed time, the format HH:MM must be used where H = hour and M = minute.

    - Example: 00:30 represents the passage of zero hours and 30 minutes.

    - When using this method to enter elapsed time, there is no need to write units.

  - Recording of Units

    - Data must always be recorded using standard unit abbreviations. The acceptable units will be consistent with the standards of the International System of Units (SI) found in the student’s textbooks.

  - Error Correction

    - All errors to CGMP documents require the following:

    - A cross-out of the error by the individual who made the error. Corrections are made by drawing a single, straight line through the entire entry containing the error.

    - A footnote represented by a number with a circle around it, such as 1 must be added to the entry line. A matching circled number is then placed at the bottom of the same page.

    - A reason for the error must be provided explaining why the original entry is no longer valid and then initialed and dated.

- Each additional correction footnote occurring on the same page as a previous footnote, must be represented by a sequential number.Attachments

N/A

| Document No.: BioWork-1070 |  | Revision No.: 02 |  | Effective Date: 23 Aug 2019 | Page 1 of 20 |
| --- | --- | --- | --- | --- | --- |
|  | Identify Errors in a Document |  |  |  |  |

# Purpose

To identify errors in a batch record. To familiarize students with the concept of Good Documentation Practices (GDP).

# Scope

This activity will focus on a Batch Record review from Johnson Community College.

# Responsibilities

It is the responsibility of students undergoing training to comply with this procedure.

# Definitions

GDP: Good Documentation Practices

Performed By / Done By / Analyzed By:	Identifies the person who actually did the work and documented it.

Verified By:	Identifies the person trained and/or qualified who ensures that the tasks performed in a step are executed and documented correctly at the time it is completed.

Approved By:	Identifies the person who evaluates the work performed to ensure it was done and/or documented accurately, completely, and in accordance with current procedures.

# References

  - Documents and Records https://www.youtube.com/watch?v=55Oj5ghLC5o

  - Good Documentation Practices (GDocP) Online Course https://www.youtube.com/watch?v=9uY1rMBP4KM

  - Alcoa Data Integrity, GDP principles https://www.youtube.com/watch?v=PFO-SUmRnRM

# Materials

N/A

# Safety

Wear lab PPE including lab coat, gloves, and safety glasses, while working in the Labs.

- Procedure (Batch Record Cover Pages)

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 2 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

Batch #:	Biotech Buffer 11172018

ISSUANCE RECORD

- This document is an accurate reproduction of the Master Record as confirmed by the preparer.

- The Batch Number appears on the first page, at a minimum.

- A Unique Identifier is affixed to every page (the Unique Identifier may be the Batch Number) as confirmed by the preparer.

- The version number is current.

- The Production Order number and/or Material number if applicable, is on the first page.

ISSUED BY: 	Bev J Owens

11 17 2018

/	/

Signature	Date

Production Order Number 	101330014		[	] N/A

Material Number 	6014300	 [	] N/A

| Legend: | Date (D) = DD/MMM/YYYY Time (t) (24 hr) = e.g., 1400 for 2:00 PM | P = Performed by V = Verified by |
| --- | --- | --- |

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 3 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

# Batch Record (Work Instructions)

General Information

| Step | Process Description | Data | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- |
| 1. | All process/WFI hoses must be rinsed prior to use according to SOP CS-000-AE-020. All pHs are performed undiluted. The pH meter must be standardized according to SOP CS-000-AN- 001040. The manway lid for the suspension vessel is to remain closed as much as possible throughout the process. Check that the equipment and work stations are clear of previous products, or materials not required for the planned process, and that equipment is clean and suitable for use. All printouts must be attached to the BPR in the appropriate location and signed and dated. The initials/date must extend across the page and the printout. |  |  |  |

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 4 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

JCC Buffer Tank Status

| Step | Process Description | Data | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- |
| 1. | Indicate which vessel is being used | [ ] 07BS145 or [ ] 07BS149 | ARC |  |
| 2. | Check that vessel being used has been cleaned within 12 hours NOTE: If vessel has not been cleaned within 12 hours, clean according to CS-JCC-AG-001 | Vessel cleaned within 12 hrs. [ ] Yes [ ] No | ARC | ARC |

JCC Dissolving Buffer Prep

| Step | Process Description | Data | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- |
| 1. | Select the process to be performed based on the process order: | [ ] 07BS145 or [ ] 07BS149 | ARC |  |
| 2. | Process Order # 101330003 JCC Paste Material # 6012940 Proceed to Step 4 to perform the JCC Buffer X process JCC Buffer X will be 25 kg. | [ ]JCC Buffer X (25kg) | N/A | N/A |
| 3. | Process # 101330014 JCC Paste Material # 6014300 Proceed to Step 25 to perform the JCC Buffer Y process JCC Buffer Y will be 30 kg | [ ]JCC Buffer Y (30kg) | ARC | JBL |
| 4. | JCC Buffer X will be 25 kg If this section is not applicable, check N/A and proceed to Step 25. | Section Applicable [] N/A | ARC |  |

| 5 | Check weigh the JCC paste weight. Material # 6012930 NOTE: Paste weight for the JCC Buffer X should be 1.0 kg. Contact supervisor is weight is not 1.0 kg. | 6012940 Paste Material Number Paste Weight is 1 kg. [ ] Yes [ ] No | JBL | ARC |
| --- | --- | --- | --- | --- |

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 5 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

| Step | Process Description | Data | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- |
| 6. | Check the expiration date for the JCC Paste. | D 1_1/_17/_18 Paste Expired [ ] Yes [ ] No | ARC |  |
| 7. | Calculate the CWFI required weight. 2_ 5_kg X 0.80 = Refer to Step 4 | 20 kg | ARC | JBL |
| 8. | Add the calculated CWFI required in Step 7 | 20 kg CWFI added | ARC |  |
| 9. | Record the time and date the CWFI was added. | T 1 0 : 3 3 D 11_/ 17_/ 18 | ARC |  |
| 10 | Take the initial pH of the JCC Buffer. Attach pH Label below. ARC 17-Nov-2018 . |  | JBL | N/A |
| 11 | Break up the JCC Buffer X paste from Step 5 and add it to the JCC Buffer | N/A | ARC | JBL |

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 6 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

| Step | Process Description | Data | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- |
| 12 | Record the time and date the paste is added to the JCC Buffer. | T 1 0 4 5 D 17 N/ ov_/2018 | ARC |  |
| 13 | Approximately every 5 minutes, mix the solution by hand for NLT 1 minute and record the pH . Continue with 5 minute time checks until the Buffer X paste is completely dissolved. Record on Mixing Sheet Below. | N/A | ARC |  |
| 14 | Record the time and date the Buffer X paste is completely dissolved. | T 1 1 1 3 D 17 N/ ov_/2018 | ARC |  |
| 15 | Calculate the elapsed time to dissolve the paste from Step 12 to Step 14 | 28 min | ARC | JBL |
| 16 | Calculate the Phosphoric Acid required. 2_ 5 _kg x 0.003 x 100 = JCC Buffer weight to be used (Step 4) | 7.5 ml | ARC | JBL |
| 17 | Measure out the amount of Phosphoric Acid listed in Step 16. Attach Chemical Withdraw Label Below. | 7.5 ml | ARC | JBL |
| 18 | Add the Phosphoric Acid from Step 17 | Phosphoric Acid Added | AR | JBL |
| 19 | Hand mix the JCC Buffer solution for NLT 2 minutes | T 1 1 4 3 D 17/No_v/ 2018 Mix Start | ARC |  |
| 20 | Mix end time | T 1 1 4 4 D 17/No_v/ 2018 Mix End | ARC |  |

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 7 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

| Step | Process Description | Data | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- |
| 21 | Calculate the elapsed time | 2 min NLT 2 min | ARC |  |
| 22 | Take a pH measurement of the JCC buffer solution; Attach pH Label NOTE: pH should be between 5.86 and 7.98. (If not contact the supervisor). | Supervisor contacted [ ] Yes [ ] No | ARC |  |
| 23 | Attach pH Label: |  |  |  |
|  | Date / Time 11/17/2018 |  |  |  |
|  | 11:53 |  |  |  |
|  | Comment 847117542 |  |  |  |
|  | U1 [1] 10.07 mS/cm |  |  |  |
|  | T1 [1] 20.2 °C ATC |  |  |  |
|  | ARC 11-17-18 |  |  |  |
| 24 | Proceed to Section Error! Reference source not found. and N/A remaining sections. | N/A | N/A | N/A |
| 25 | JCC Buffer Y will be 30 kg. If this section is not applicable, check N/A and proceed to Buffer Final Formulation section. | Section Applicable [ ] N/A |  |  |
| 26 | Check weigh the JCC paste weight. |  |  |  |
|  | Material # 6014200 |  |  |  |
|  | NOTE: Paste weight for the JCC Buffer Y should be 1.5 kg. Contact supervisor is weight is not 1.5 kg. | Paste Material Number |  |  |
|  |  | Paste Weight is 1.5 kg. [ ] Yes [ ] No |  |  |
| 27 | Check the expiration date for the JCC Paste | D / / Paste Expired [ ] Yes [ ] No |  |  |

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 9 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

| Step | Process Description | Data | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- |
| 28 | Calculate the CWFI required weight. kg X 0.80 = Refer to Step 25 | kg |  |  |
| 29 | Add the calculated CWFI required in Step 28 | kg CWFI added |  |  |
| 30 | Record the time and date the CWFI was added. | T D / / |  |  |
| 31 | Take the initial pH of the JCC Buffer. Attach pH Label below. |  |  |  |
| 32 | Attach pH Label: |  |  |  |
| 33 | Break up the JCC Buffer Y paste from Step 26 and add it to the JCC Buffer. |  |  |  |
| 34 | Record the time and date the paste is added to the JCC Buffer. | T D / / |  |  |

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 10 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

| Step | Process Description | Data | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- |
| 35 | Approximately every 5 minutes, mix the solution by hand for NLT 1 minute and record the pH. Continue with 5 minute time checks until the Buffer Y paste is completely dissolved |  |  |  |
| 36 | Record the time and date that Buffer Y paste is completely dissolved. | T D / / |  |  |
| 37 | Calculate the elapsed time to dissolve the paste from Step 34 to Step 36. | min |  |  |
| 38 | Calculate the Phosphoric Acid required. kg x 0.003 x 100 = JCC Buffer weight to be used (Step 25) | ml |  |  |
| 39 | Measure out the amount of Phosphoric Acid listed in Step 38 | ml |  |  |
| 40 | Add the Phosphoric Acid from Step 39 | Phosphoric Acid Added |  |  |
| 41 | Hand mix the JCC Buffer solution for NLT 2 minutes. | T D / / Mix Start |  |  |
| 42 | Mix end time. | T D / / Mix End |  |  |

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 11 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

| Step | Process Description | Data | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- |
| 43 | Calculate the elapsed time | min NLT 2 min |  |  |
| 44 | Take a pH measurement of the JCC buffer solution. NOTE: pH should be between 5.86 and 7.98. (If not contact the supervisor). | Supervisor contacted [ ] Yes [ ] No |  |  |
| 45 | Attach pH Label: |  |  |  |

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 12 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

JCC Buffer Final Formulation

| Step | Process Description | Data | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- |
| 1. | 25 kg + 20 kg + 1_.0 kg = Step 4 Step 8 Step 5 or or or Step 25 Step 29 Step 26 | 1 4 . 0 3. 0 kg | ARC | JBL |
| 2. | Add the calculated CWFI required in Section Error! Reference source not found.. | 4_..0 kg | TST | JBL |
| 3. | Hand mix the JCC Buffer solution for NLT 2 minutes. | T 1 2 0 1 D 17 / No/v 18 Mix Start | TST |  |
| 4. | Mix end time. | T 1 2 0 3 D 17 / No/v 18 Mix End | TST |  |
| 5. | Calculate the elapsed time. | 2 min NLT 2 min | TST |  |
| 6. | Check the conductivity of the JCC Buffer. Attach Conductivity print-out below. | 1 0 . 18 mS/cm² |  |  |
| 7. | Date / Time 11/17/2018 12:11 Comment 847117542 U1 [1] 10.16 mS/cm T1 [1] 20.0 °C ATC ARC 17-Nov-18 | N/A | ARC |  |

.

| Notes: 1 Calculation erro . ARC 11-17-18 |  | . |
| --- | --- | --- |
|  |  | BioWork Lab |

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 13 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

JCC Buffer Final Formulation (Continued)

| Step | Process Description | Data | Performed By / Date: |  | Verified By / Date: |
| --- | --- | --- | --- | --- | --- |
| 8. | Calculate the amount of 1 M NaOH to add to the buffer. 2 5 kg x 5.018 mL/kg = Step 4 or 25 | 1 2 6 ml | ARC |  | JBL |
| 9. | Record the 1 N NaOH buffer number and expiration date. | 1 N NaOH Buffer # 1 25-Dec-26 D31/ No_/v 18 Expiration Date |  | ARC | JBL |
| 10 | Measure out the amount of NaOH buffer calculated in Step 8. | 1 2 _6_ml | ARC |  | JBL |
| 11 | Attach Chemical Withdrawal Label to Batch Record | N/A | JBL |  |  |
| 12 | Dispense the NaOH buffer measured out in Step 10 into the JCC Buffer. |  |  |  | JBL |
| 13 | Mix the JCC Buffer for NLT 2 minutes. | T 1 2 3 0_ D17/Nov/ 18 Mix start | JBL |  |  |
| 14 | Solution mix end time. | T 1 2 4 9_ D17/No_v/ 18 Mix End | JBL |  |  |
| 15 | Calculate the elapsed time for the JCC Buffer mix. | 19 min NLT 2 min | JBL |  |  |

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 14 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

Filtration Of The JCC Buffer

| Step | Process Description | Data | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- |
| 1. | Obtain a 10 inch 4.5 µm encapsulated filter. Record the batch number and expiration date. | 7008502 Batch # D30/Jun/ 2_7 Expiration Date | JBL |  |
| 2. | Make the following connections: - Connect a hose from the bottom of the JCC Buffer vessel to a transfer pump. - Connect a hose from the transfer pump to the filter listed in Step1. - Connect a hose from the filter to transfer vessel 07BS143 (check vessel to insure the Clean Hold Time has not expired). | 07BS143 Transfer Vessel ID Clean Hold Time expired [ ] Yes [ ] No | JBL | ARC |
| 3. | Start the transfer of the JCC Buffer into vessel 07BS143. Record the start time and date. | T 1 3 0 0 D17/_No/v 1_8 Filtration Start | JBL |  |
| 4. | Filtration end time and date. | T 1 4 2 2 D 17/N o_v/ 18 | JBL |  |
| 5. | Calculate the elapsed time for the JCC Buffer filtration. | 82 min | JBL |  |

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 15 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

| Step | Process Description | Data | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- |
| 6. | Calculate the expiration date of the JCC Buffer. Expiration is 12 days and 6 hours from the filtration end time and date Step 4. | Filtration End Step 4. D22/Nov/18 T 1 4 2 2 Expiration 12 days 6 Hours D 04/Dec/ 18 T 2 0 2 2 | JBL | ARC |
| 7. |  |  |  |  |

The following signature signifies that all sections of this document have been reviewed and completed.

/	/

Supervisor/Designate

Date

endbpr

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 16 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

# Attachments

Buffer X Mixing Table	(JCC Buffer Y will be 25 kg)

| Time | Date | Mixed 1 min | pH | Initial |
| --- | --- | --- | --- | --- |
| t 1 _0_ 5 5 | D 11/ 1_7/ | [ ] Yes [ ] No | 6 ._9 9 | ARC |
| t 1 _1 _ 0 0 | D 1/1 1/7 | [ ] Yes [ ] No | 6 ._9 9 | ARC |
| t 1 _1_ _0_ _5_ | D 1/1 1_7/ | [ ] Yes [ ] No | 7._0 0 | ARC |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 17 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

Buffer X Mixing Table	(JCC Buffer Y will be 30 kg)

| Time | Date | Mixed 1 min | pH | Initial |
| --- | --- | --- | --- | --- |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |
| t | D / / | [ ] Yes [ ] No | . |  |

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 18 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

ATTACH CHEMICAL LABEL

__________

Chemical Withdrawal Label	(JCC Buffer Y will be 25 kg)

| ATTACH CHEMICAL LABEL Chemical Withdrawal Label Chemical Name: Phosphoric Acid SAP Material #: 70089187 Batch #: 27BA9001807 Expiration: 11-07-18 Weighed by: JBL Verified by: ARC Attach Weigh Label Here |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 19 of 20 |
|  | Identify Errors in a Document |  |  |  |

Chemical Withdrawal Label	(JCC Buffer Y will be 30 kg)

| Notes: |  |
| --- | --- |
|  | BioWork Lab |

ATTACH CHEMICAL LABEL

Chemical Withdrawal Label

Chemical Name: 			 SAP Material #: 	 Batch #: 	 Expiration: 		 Weighed by: 		 Verified by:

Attach Weigh Label Here

| Document No.: BioWork-1070 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 20 of 20 |
| --- | --- | --- | --- |
| Identify Errors in a Document |  |  |  |

Chemical Withdrawal Label	Buffer Final Formulation (JCC Buffer Y will be 25 kg)

ATTACAHTCTAHCEMH ICHAELMLIACBAELL

Chemical Withdrawal Label

Chemical Name:  NaOH		 SAP Material #:   27PL486  Batch #:  89252600081 Expiration:  25-Dec-2026	 Weighed by: 	JBL	 Verified by:   ARC

Attach Weigh Label Here

126 mL

JBL / ARC

| Document No.: BioWork-1080 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 1 of 5 |
| --- | --- | --- | --- |
| Write a SOP – Making a Peanut Butter and Jelly Sandwich |  |  |  |

#### Purpose

The purpose is to experience the process of writing a standard operating procedure for a familiar activity - making a peanut butter and jelly sandwich.

#### Scope

As a team, fill out this SOP template with all of the information needed to document the process of making a peanut butter and jelly sandwich.

#### Responsibilities

It is the responsibility of students undergoing training to comply with this procedure. Authors of Documents are responsible for the following:

- Creating a new Standard Operating Procedure according to the requirements as stated in this SOP, ensuring that the procedure is detailed, succinct and accurate.

- Adhering to content and format requirements and the use of established document templates to generate procedures.

- Resolving content issues with Reviewers and once all issues are resolved, submitting the procedure to the Document Coordinator for processing.

- Ensuring that all forms related to a particular SOP are reviewed for content and accuracy each time an SOP is revised.

Document Reviewers are responsible for verifying correct content and formatting of all submitted documents and facilitating the document approval and issuance process.

#### Definitions

Any definitions here

#### References

Any references here

| Document No.: BioWork-1080 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 2 of 5 |
| --- | --- | --- | --- |

| Notes: |  |
| --- | --- |
|  | BioWork Lab |

Write a SOP – Making a Peanut Butter and Jelly Sandwich

#### Materials

Any materials here

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |

#### Safety

All safety information here.

#### Procedure

Example

  - Stepwise procedure steps here.

    - These are substeps

      - This is a sub substep

Important:	this is important info.

| Document No.: BioWork-1080 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 3 of 5 |
| --- | --- | --- | --- |
| Write a SOP – Making a Peanut Butter and Jelly Sandwich |  |  |  |

| Procedure (For Making a Peanut Butter & Jelly Sandwich)Document No.: BioWork-1080 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 4 of 5 |
| --- | --- | --- | --- |
| Write a SOP – Making a Peanut Butter and Jelly Sandwich |  |  |  |

#### Procedure (PB&J continued)

| Document No.: BioWork-1080 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 5 of 5 |
| --- | --- | --- | --- |
| Write a SOP – Making a Peanut Butter and Jelly Sandwich |  |  |  |

- Figures and Diagrams

Diagram 1: (If needed, any diagram or drawing)

#### Notes/Observations

- Attachments

N/A

| Document No.: BioWork-2010 | Revision No.: 01 | Effective Date: 23 Aug 2019 | Page 1 of 4 |
| --- | --- | --- | --- |
| Lab Orientation |  |  |  |

#### Purpose

  - To tour the lab to become familiar with safety features common to bioprocessing, chemical, and pharmaceutical manufacturing workplaces.

#### Scope

  - As your instructor takes you through the teaching laboratory, note safety features and safety equipment. This activity has six parts:

    - Part 1: Lab Rules

    - Part 2: Lab Equipment

    - Part 3: Lab Glassware

    - Part 4: Using Glassware and Reading Glassware Scales

    - Part 5: Lab Organization, Clean-up, and Storage

    - Part 6: Working in Groups

#### Responsibilities

  - Technicians will operate equipment under the direction of an instructor.

#### Definitions

  - N/A

#### References

  - N/A

#### Materials

  - graduated cylinders of two different sizes

  - pipettes of two different sizes

  - pipette aid

  - beaker (500 mL)

#### Safety

  - Wear lab PPE including lab coat, gloves, and safety glasses.

#### Procedure

  - Lab Rules

    - Wear lab coat and safety glasses at all times when in the lab.

    - Never smoke, drink, or eat when in the lab.

    - Keep personal items stored in the proper location.

    - If an accident occurs, inform your instructor immediately.

    - When in a new work area or lab, make it your first step to become familiar with the location and use of all safety features, equipment, and procedures.

    - Use the correct type of water. As a process technician, you will rarely use ordinary tap water—because most processes require distilled or deionized water. In this course:

| Document No.: BioWork-2010 | Revision No.: 01 | Effective Date: 23 Aug 2019 | Page 2 of 4 |
| --- | --- | --- | --- |
| Lab Orientation |  |  |  |

      - When DI water is listed as a material for a lab activity, it means distilled or deionized water.

      - When water is listed as a material for a lab activity, it means tap water.

#### Lab Equipment

    - In addition to safety equipment, you need to become familiar with basic equipment on the chart that follows. You will learn more about these and other items during lab activities in the course.

| Equipment | Purpose |
| --- | --- |
| content of workbench drawers | holds supplies for lab work |
| sink | provides water supply and place to wash |
| refrigerator | stores perishable supplies |
| autoclave | sterilizes materials |
| dishwasher | washes glassware |
| incubator | grows cell cultures |
| water bath | maintains materials at constant temperature |
| balance | weighs materials |
| pH meter | measures acidity or alkalinity of samples |
| microscope | makes it possible to see microorganisms by magnifying them |
| Magnetic stirrer/hot plate | stirs and/or heats liquids |

#### Lab Glassware

    - You also need to become familiar with various glassware used in the lab—and why and when it is used. The chart that follows describes basic lab glassware.

| Glassware | Description (All glassware comes in various sizes.) |
| --- | --- |
| beaker | holds materials measures volume approximately |
| graduated cylinder | measures volume of liquids precisely |
| flask | holds liquids triangular flask is called an Erlenmeyer flask |
| bottle | holds liquids for sterilizing or storage |
| pipette | transfers and measures liquids looks like a glass or plastic straw |
| pipette aid | device that attaches to pipette to draw in and release liquid |

#### Using Glassware and Reading Glassware Scales

    - Collect these items from your instructor.

      - graduated cylinders of two different sizes

      - pipettes of two different sizes

      - pipette aid

      - beaker (500 mL)

| Document No.: BioWork-2010 | Revision No.: 01 | Effective Date: 23 Aug 2019 | Page 3 of 4 |
| --- | --- | --- | --- |
| Lab Orientation |  |  |  |

    - Note that the graduated cylinders and pipettes have graduations, or marks, in a scale. The scales indicate measurements in milliliters (mL) of volume. Later in this unit you will learn about milliliters. For now, just concentrate on how to read the scales on the glassware.

    - Put some tap water into the smaller of the graduated cylinders.

    - Hold the cylinder so the water line is at eye level. You will see something like the picture below. The water level appears curved. This curved liquid surface is called a meniscus. Always read the scale at the bottom of the meniscus. In the example shown below, the correct reading is 20.

    - Note: that the curve of the meniscus flattens in wider cylinders and

those made of plastic.

    - Fill one of the graduated cylinders about half full with tap water. It is not important exactly how much water you add. Now read the scale. How many milliliters of water do you have?

    - Put some tap water in your other graduated cylinder. How many milliliters of water do you have?

    - Put about 200 mL of tap water in your beaker.

    - Practice attaching the pipette aid to the large pipette. Now draw water into the pipette from the beaker and then expel the water from the pipette.

    - Now fill the large pipette approximately half-full. Read the scale on the pipette. Notice that the meniscus is more curved in the pipette than in the graduated cylinder.

      - How many milliliters of water do you have?

    - Repeat the above step with the small pipette.

      - How many milliliters of water do you have?

#### Lab Organization, Clean-up, and Storage

    - Now you are familiar with the locations of many items that you will use in course activities. Listen and observe as your instructor describes and shows you:

      - Where to find materials for use in specific activities.

      - How to wash dirty glassware.

      - Where to place dirty glassware that you are not responsible for washing at the end of an activity.

      - Where to dispose of used pipettes, gloves, and other disposable items.

| Document No.: BioWork-2010 | Revision No.: 01 | Effective Date: 23 Aug 2019 | Page 4 of 4 |
| --- | --- | --- | --- |
| Lab Orientation |  |  |  |

      - How to clean your bench top at the end of an activity.

      - Where to store your personal items during work in the lab.

#### Working in Groups

    - The ability to effectively communicate and work as a team is an invaluable skill for employee success. One way to learn team skills is to practice different roles (described below) as you complete activities throughout this course.

    - The more roles you practice, the more confident and effective you will become as a team member—and by assigning different roles during every activity, your group work will proceed more smoothly.

    - Listen as your instructor explains each role in the checklist.

| Group Role Checklist |  |
| --- | --- |
| Supplier | collects materials needed |
| Reader | reads instructions carefully and explains them to the operator |
| Operator | carries out the steps in the procedure (measuring, pouring, mixing, etc.) |
| Recorder | records data collected in the activity |
| Verifier | double-checks supplies, instructions, measurements, data, etc. |

- Attachments

  - N/A

| Document No.: BioWork-2020 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 1 of 5 |
| --- | --- | --- | --- |
| PPE Exhibit & Scenarios |  |  |  |

#### Purpose

To learn about the kinds of PPE you may need on the job in various settings. To experience what it feels like to wear or use several types of PPE.

#### Scope

As your instructor takes you through the teaching laboratory, note:

- The various types of PPE for different activities.

- Your comfort level with the PPE.

#### Responsibilities

Technicians will operate equipment under the direction of an instructor. It is the responsibility of students undergoing training to comply with this procedure.

#### Definitions

PPE: Personal Protective Equipment

SCBA:Self-contained Breathing Apparatus

#### References

- N/A

#### Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Safety glasses | 1 | Each |
| Safety shoes |  |  |
| Surgical mask |  |  |
| Hair cover |  |  |
| Beard cover |  |  |
| Latex gloves |  |  |

| Disposable overalls |  |  |
| --- | --- | --- |

| Document No.: BioWork-2020 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 2 of 5 |
| --- | --- | --- | --- |
| PPE Exhibit & Scenarios |  |  |  |

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Disposable overshoes |  |  |
| Full face forced air respirator or SCBA |  |  |
| PVC gloves |  |  |
| PVC suit |  |  |
| PVC boots |  |  |
| Hard hat |  |  |
| Work gloves |  |  |
| Lifeline and harness |  |  |
| Goggles |  |  |
| Hard hat with attached face shield |  |  |
| Rubber gauntlets |  |  |
| Rubber apron |  |  |
| Ear plugs |  |  |

- Safety

Wear lab PPE including lab coat, gloves, and safety glasses

#### Procedure

  - General

    - Listen as your instructors shows you a variety of PPE items displayed on a table.

    - Identify the correct PPE for each of the of the following scenarios. Have one student from your team dress appropriately for the scenario.

| Document No.: BioWork-2020 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 3 of 5 |
| --- | --- | --- | --- |
| PPE Exhibit & Scenarios |  |  |  |

  - Scenario 1

    - A media-mixing tank in your section of the plant is two stories high. On the last shift, a blade on the shaft of the tank agitator worked loose. This agitator mixes the contents of the tank. In preparation for the repair job, you and your coworkers flush the tank with water, then lock out, and tag the agitator motor to prevent it from being turned on. Then you check the atmosphere inside the tank and you see that it is safe. To repair the agitator, a maintenance technician must enter the tank through a hatch at the top. The tank has no forced ventilation. The maintenance technician asks you to bring him the correct PPE.

    - What PPE is needed?

  - Scenario 2

    - You are assigned to the co-mill in the granulating department of a pharmaceutical company. The mill is somewhat like a giant blender with many blades and it is very noisy. At an early step in the process, the mill mixes ingredients that will later be formed into tablets for medicine. Today, Product X is ready to be put in the co-mill. It is moderately toxic in large quantities. Product X has been drying in a big bowl. There are about 200 lb of it open to the room. The bowl of Product X is attached to a bowl hoist and lifted over the co-mill. A lot of dust is created when Product X is poured into the mill and a local exhaust fan picks up most of the dust.

    - What PPE is needed?

| Document No.: BioWork-2020 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 4 of 5 |
| --- | --- | --- | --- |
| PPE Exhibit & Scenarios |  |  |  |

  - Scenario 3

    - Any foreign matter sucked into a pump could damage it, so some pumps have strainers. On a 30% hydrochloric acid piping system in your area, a strainer is clogged and you are assigned to clean it. The line has been flushed with water, but there may be some residual acid present. Even though the system pressure reads zero, there may be slight pressure. When the strainer housing is unbolted, this pressure may cause any residual chemical to spray out. Hydrochloric acid is corrosive and can cause severe burns. Inhaling the vapors is dangerous.

    - What PPE is needed?

  - Scenario 4

    - You enter the chemical storage area of the plant on a routine walk-through and find that a 55-gallon drum of caustic soda (25%) has fallen off a broken pallet and spilled over a 50-foot-wide area. Caustic soda is corrosive and can cause skin irritation and burns. Inhaling the vapors can cause serious damage to the upper respiratory tract. You phone the plant emergency response team leader and she instructs you to rope off the area until the team can arrive at the scene. Because you have been trained as a member of the plant’s emergency response team, the leader asks for your assistance in cleaning up the spill.

    - What PPE is needed?

| Document No.: BioWork-2020 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 5 of 5 |
| --- | --- | --- | --- |
| PPE Exhibit & Scenarios |  |  |  |

  - Scenario 5

    - In the preparation of fermentation medium, all liquid raw materials are added automatically to the media-mixing tanks through a computer-controlled system. Your job is to add all dry raw materials. These materials are bagged or drummed. The process technician climbs to a platform with handrails at the top of the tank and from the platform adds the materials through a hopper. These dry raw materials include sucrose and sodium chloride, also known as sugar and salt. Maximum container weight for these materials is 50 pounds. Large quantities can create nuisance dust, which can irritate the eyes and respiratory tract.

    - What PPE is needed

  - Scenario 6

    - You work in the cell culture area of a biotechnology company. Your task is to monitor the growth of a disease-causing bacterium in 3,000-gallon bioreactors.

The organism is grown as the first step in producing a vaccine against it. You must protect yourself from contact with the organism and prevent any bacterial cells from accidentally getting out of the production area. Entry and exit into the production area are strictly controlled.

    - What PPE is needed?

#### Attachments

N/A

| Document No.: BioWork-2030 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 1 of 3 |
| --- | --- | --- | --- |
| J.J. Keller’s OSHA Safety Training Handbook |  |  |  |

#### Purpose

To become familiar with the J.J. Keller’s OSHA Safety Training Handbook.

To practice quickly locating and using job-related information, including documentation and reference materials To practice making decisions and working toward a common goal with others.

#### Scope

You will work with your team to locate information in the J.J. Keller’s OSHA Safety Training Handbook and answer specific questions. Then you will participate in a follow-up discussion with the class.

#### Responsibilities

It is the responsibility of students undergoing training to comply with this procedure. Technicians will operate equipment under the direction of an instructor.

#### Definitions

OSHA:	Occupational Safety and Health Administration

#### References

- J.J. Keller’s OSHA Safety Training Handbook; 7th or 8th edition

#### Materials

N/A

#### Safety

Wear lab PPE including lab coat, gloves, and safety glasses.

#### Procedure

  - In the J.J. Keller’s OSHA Safety Training Handbook, read the introductory page entitled “Introduction to the Worker”.

  - Your instructor will assign your team two reviews from the J.J. Keller’s OSHA Safety Training Handbook. There is a review at the end of each of the 25 chapters. See list of Chapters (topics)

  - Work with your team to complete the two reviews and enter answers in the appropriate column on the activity form.

| Document No.: BioWork-2030 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 2 of 3 |
| --- | --- | --- | --- |
| J.J. Keller’s OSHA Safety Training Handbook |  |  |  |

  - Decide as a group how to quickly locate the correct answers in the time available and complete as many questions as you can.

  - If you think you can move faster by trying a different team approach, do so. Remember to make good use of the table of contents, headings, and illustrations to locate information.Step 2

#### Notes/Observations

Topic Index

|  | Topic | Page |
| --- | --- | --- |
| 1 | An Introduction to OSHA | 1 |
| 2 | Company Vehicle Safety | 11 |
| 3 | Confined Space Entry | 23 |
| 4 | Electrical Safety | 31 |
| 5 | Emergency Response | 45 |
| 6 | Eye Protection | 55 |
| 7 | Fall Protection | 63 |
| 8 | Fire Prevention | 73 |
| 9 | First Aid and Bloodborne Pathogens | 85 |
| 10 | Foot Protection | 101 |
| 11 | Forklift Safety | 109 |
| 12 | Hand Protection | 119 |
| 13 | Hazard Communication | 125 |
| 14 | Head Protection | 141 |
| 15 | Hearing Conservation | 149 |
| 16 | Lifting and Ergonomics | 159 |
| 17 | Lockout/Tagout | 173 |
| 18 | Machine Guarding | 183 |
| 19 | Materials Handling Equipment | 193 |
| 20 | Respiratory Protection | 203 |
| 21 | Slips, Trips and Falls | 213 |
| 22 | Tool Safety | 225 |
| 23 | Violence in the Workplace | 237 |
| 24 | Welding, Cutting, and Brazing | 243 |
| 25 | Workplace Security | 257 |

#### Attachments

N/A

| Document No.: BioWork-2030 | Revision No.: 02 | Effective Date: 23 Aug 2019 | Page 3 of 3 |
| --- | --- | --- | --- |
| J.J. Keller’s OSHA Safety Training Handbook |  |  |  |

#### Activity Form (for 8th Ed)

|  | Lockout/ Tagout | Eye Protection | Hazard Communication | Lifting & Ergonomics |
| --- | --- | --- | --- | --- |
| 1 |  |  |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
| 4 |  |  |  |  |
| 5 |  |  |  |  |
| 6 |  |  |  |  |
| 7 |  |  |  |  |
| 8 |  |  |  |  |
| 9 |  |  |  |  |
| 10 |  |  |  |  |

Note: Key hidden

#### Activity Form (for 7th Ed)

|  | Lockout/ Tagout | Eye Protection | Hazard Communication | Lifting & Ergonomics |
| --- | --- | --- | --- | --- |
| 1 |  |  |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
| 4 |  |  |  |  |
| 5 |  |  |  |  |
| 6 |  |  |  |  |
| 7 |  |  |  |  |
| 8 |  |  |  |  |
| 9 |  |  |  |  |
| 10 |  |  |  |  |

Note: Key hidden

| Document No.: BioWork-2050 | Revision No.: 01 | Effective Date: 23 Aug 2019 | Page 1 of 4 |
| --- | --- | --- | --- |
| Safety Survey |  |  |  |

#### Purpose

To provide a training exercise to teach and reinforce the beneficial exercise of conducting regular Safety Surveys of all functional areas throughout the facility.

#### Scope

Training should be provided to everyone supervising or working in a given area to ensure that safety survey participants understand and apply the required approach to get the most effective outcome from the survey process.

#### Responsibilities

Trainees will be lead through the exercise under the direction of an instructor. It is the responsibility of students undergoing training to comply with this procedure.

The Instructor will set up a sample area where the Safety Survey will take place. This area will be “seeded” with various safety concerns and/or violations that the trainees should see, recognize as hazards and record on the Safety Survey Form. Examples of how a given area might be set up are provided in the Section 8.

Trainees are responsible for completing a Safety Survey walk through the identified area and then completing a Safety Survey Checklist / Report for the given scenario.

The Instructor will then review the results of the Safety Survey exercise and provide feedback as appropriate.

#### Definitions

Safety Survey:	A tool used by supervisors and employees to proactively identify potential hazards and to eliminate or control those hazards before they result in an incident and/or injury. Through a critical examination of the workplace, safety surveys help to identify and record hazards for corrective action.

Regular workplace safety surveys are an important part of the overall occupational health and safety program and safety management system.

#### References

- Canadian Centre for Occupational Health and Safety – “Effective Workplace Inspections”

#### Materials

- Process Area Safety Survey Checklist example (Attachment 9.1)

- Laboratory Safety Survey Checklist example (Attachment 9.2)

- Various props to be placed in the area to be surveyed

#### Safety

| Document No.: BioWork-2050 | Revision No.: 01 | Effective Date: 23 Aug 2019 | Page 2 of 4 |
| --- | --- | --- | --- |
| Safety Survey |  |  |  |

Instructor is to provide all required personal protective equipment and ensures that it is worn by all trainees before entering the process area/laboratory where the training exercise will take place.

#### Procedure

  - Safety Surveys are important as they allow you to:

- Listen to the concerns of workers and supervisors

- Main further understanding of jobs and tasks

- Identify existing and potential hazards

- Determine underlying causes of hazards

- Recommend corrective action

- Monitor steps taken to eliminate hazards or control the risk (e.g., engineering controls, administrative controls, policies, procedures, personal protective equipment.

- Safety Surveys should only be done when there are people working in the area to be surveyed. This is because one of the most important elements of a successful survey program is for the survey team to speak with the people who actually work in the area – the people who are the most familiar with the hazards in that space as they work in the area every day.

  - The instructor determines what type of area will be used to conduct the safety survey training (e.g. process area vs. laboratory).

  - The instructor sets up the area to be surveyed by “planting” any number of items or scenarios that should be noted and addressed during a safety survey.

    - The instructor can make use of the area checklists for ideas on what can be set up (see attachments 9.1 or 9.2).

    - If at all possible, either the instructor or another person should play the role of an employee that works in the area to be surveyed. It would only be a successful survey if at least one of the survey team members interviews the area worker.

  - Possible items to set up in an area to be surveyed include:

- Place a cart or run an electrical wire across an exit or in front of an eyewash station.

- Block the area walkways or access to the eyewash station with chairs or carts.

- Ladders or step stools are not marked with an up to date inspection sticker

- One or more of the area lights can be “burned out” by unscrewing the light or removing it.

- Remove the area Evacuation Plan posting

- No posting for required personal protective equipment.

- Set up an eyewash with an out of date inspection tag

- Have the area worker wearing inappropriate footwear (like sneakers or sandles).

- If there is a chemical fume hood or other administrative control in the area, place an inspection tag on it that shows it is overdue for service.

- Place a food item in the area.

- Drap a long extension cord so that it dangles in an isle or presents a trip hazard

Page 3 of 4

Effective Date:

23 Aug 2019

Revision No.:

01

Place containers of “chemicals” in the area that do not have proper (or any) labels. Document NoN.:otes:

## BioWork-2050

- Place two bottle of water – one labeled as an acid and one label

Safety Survey

in a storage area with no secondary containment.

ed as a baBsieoWneoxrtktoLaebach other

- Place an out of date inspection tag on the area fire extinguisher

- Place a small puddle of water on the floor at the base of a piece of equipment

- Place a piece of equipment on the bench that has a frayed or damaged cord.

  - Depending on the number of people in the training exercise, the instructor may assign the following roles to trainees for purposes of this training exercise. The instructor will determine how many of each “role” will be assigned to best fit the scope of the training exercise.

    - An employee that works in the area being surveyed and is familiar with the operations and safety concerns of that area

    - Supervisor of the area where the survey is being done. This person should lead the Safety Survey team.

    - Safety Department representative – that makes sure the Safety Survey team members remain engaging and interactive throughout the survey.

  - The Instructor then provides each trainee with a copy of the appropriate Safety Survey checklist and reminds them that the checklist are just a tool and not to be the sole focus of the survey.

  - The trainees then enter the area to be surveyed and complete the Safety Survey.

  - The Instructor accompanies the Safety Survey team and should provide real time, appropriate feedback as needed. Things for the instructor be cognizant include:

- Be sure everyone on the Safety Survey team remain safe at all times.

- The Survey team members don’t become overly dependent on the checklist tool. It is common for inexperienced Safety Survey team members to turn the exercise into a simple “checkbox” exercise and move through the area quickly as if they were just checking boxes and inspecting the area. Avoid this at all cost. When Safety Surveys are done like this, they are often very ineffective and most importantly, leave the people working in the area feeling like they are being scrutinized unfairly. The exact opposite effect the Safety Survey should have.

- Be sure that the team is working together to cover the space effectively and to interact with as many of the area workers as possible. Good Safety Surveys are “less inspection and more conversation” in nature.

- Be sure the team members are taking notes and recording their finds, questions, or comments from area workers.

  - Once the Safety Survey team indicates that they have completed their walkthrough but before anyone is allowed to leave the area, review the various checklists and notes to confirm that all of the “set up” items that were placed in the area have been found and documented. Review any of those that may have been missed.

| Document No.: BioWork-2050 | Revision No.: 01 | Effective Date: 23 Aug 2019 | Page 4 of 4 |
| --- | --- | --- | --- |
| Safety Survey |  |  |  |

  - After leaving the area that was just surveyed, the group should sit down and compare notes to develop a summary report of all the findings.

    - The report should look very much like an Incident Analysis form in that it should contain the following sections:

      - Finding - 1 sentence summary of what was found

      - Corrective Actions – The team should work together to identify what potential corrective action(s) will be put into place to correct each finding. It should be noted that not all Findings are created equally. Some findings are easy to fix immediately or very quickly at very little cost while on the other end of the spectrum there may be findings that require corrective actions that may take time and a significant expense to implement. These corrective actions should take the form of a request to upper management.

      - Person Responsible – For each Corrective Action, a person should be assigned to make certain the action is taken in a timely fashion (see Date to be Completed next)

      - Date to be Completed – Each Corrective Action must include a proposed completion date. This date should be based on factors including how great the hazard is, how likely it will be for someone to be exposed to the hazard and the complexity and expense of implementing the action.

      - Date Completed – This is a great tool to measure the effectiveness of the Company Safety Management system. Identified Corrective Actions that were completed on time can be used as a meaningful metric for overall safety performance.

  - The Instructor then communicates to the trainees how important it is to complete this Safety Survey report and submit it to department management and the Safety department in a timely manner.

  - Instructor then asks if there are any final questions from the trainees and addresses them accordingly.

#### Attachments

  - Process Area Safety Survey Checklist example (Activity 2050.1)

  - Laboratory Safety Survey Checklist example (Activity 2050.2)

| Document No.: BioWork-2060 | Revision No.: 01 | Effective Date: 20 Jul 2020 | Page 1 of 13 |
| --- | --- | --- | --- |
| Lockout Tagout Procedure & Demonstration |  |  |  |

#### Purpose

To provide an effective “hands on” training exercise to review the Control of Hazardous Energy Sources through Lockout / Tagout procedures..

#### Scope

A Company is required to provide training that must ensure that employees understand the purpose, function, and restrictions of the Company energy-control program. Employers must provide training specific to the needs of “Authorized”, “Affected”, and “Other” employees.

#### Responsibilities

It is the responsibility of students undergoing training to comply with this procedure. Trainees will operate the lockout/tagout equipment and any associated equipment under the direction of an instructor.

  - Your employer shall have the responsibility to provide training to ensure that the purpose and function of the energy control program are understood by employees and that the knowledge and skills required for the safe application, usage, and removal of the energy controls are acquired by employees.

  - Instructor will assign the following roles to trainees for purposes of this training exercise. The instructor will determine how many of each “role” will be assigned to best fit the scope of the training exercise. See Section 4 for the definitions of “Authorized Employee”, “Affected Employees” and “Other Employees”.

- System Owner – Person(s) responsible for the area/department/operations of the equipment impacted by the Lockout/Tagout process.

- Authorized Employees

- Affected Employees

- Other Employees

#### Definitions

“Affected Employees”	(usually machine operators or users) are employees who operate the relevant machinery or whose jobs require them to be in the area where service or maintenance is performed. These employees do not service or maintain machinery or perform lockout/tagout activities. Affected employees must also receive training in the purpose and use of energy control procedures. They also need to be able to do the following:

- Recognize when an energy control procedure is being used

- Understand the purpose of the procedure

- Understand the importance of not tampering with lockout or tagout devices and NOT starting or using equipment that has been locked and/or tagged out.

| Document No.: BioWork-2060 | Revision No.: 01 | Effective Date: 20 Jul 2020 | Page 2 of 13 |
| --- | --- | --- | --- |
| Lockout Tagout Procedure & Demonstration |  |  |  |

“Authorized Employees”	are those with responsibilities to implement the energy control procedures or performing the service and/or maintenance activities. They need the knowledge and skills necessary for the safe application, use, and removal of energy-isolated devices. They must also have training in the following:

- Hazardous energy source recognition

- The type and magnitude of the hazardous energy sources in the workplace; and

- Energy-control procedures, including the methods and means to isolate and control those energy sources.

Capable of being locked out.	An energy isolating device is capable of being locked out if it has a hasp or other means of attachment to which, or through which, a lock can be affixed, or it has a locking mechanism built into it. Other energy isolating devices are capable of being locked out, if lockout can be achieved without the need to dismantle, rebuild, or replace the energy isolating device or permanently alter its energy control capability.

Energized	Connected to an energy source or containing residual or stored energy.

Energy isolating device A mechanical device that physically prevents the transmission or release of energy.

Energy source	any source of electrical, mechanical, hydraulic, pneumatic, chemical, thermal or other energy. Examples are:

- Energized electrical parts

- Hydraulic or pneumatic pressure

- Pressurized pipes or vessels

- Flywheels

- Batteries

- Capacitors

- Thermal energy (e.g. steam, heated liquids)

- Residual chemicals causing thermal or pressure build-up

- Gravity

- Static energy

Hazardous Energy Control Program (HECP) The Company written program that provides specific guidance on how to identify hazardous energy sources and what specific procedures must be followed to neutralize and isolated said energy sources as well as procedures that must be followed to safely reenergized any equipment.

Lockout The placement of a lockout device on an energy isolating device, in accordance with an established procedure, ensuring that the energy isolating device and the equipment being controlled cannot be operated until the lockout device is removed.

| Document No.: BioWork-2060 | Revision No.: 01 | Effective Date: 20 Jul 2020 | Page 3 of 13 |
| --- | --- | --- | --- |
| Lockout Tagout Procedure & Demonstration |  |  |  |

Lockout device A device that utilizes a positive means such as a lock, either key or combination type, to hold an energy isolating device in the safe position and prevent the energizing of a machine or equipment.

“Other Employees”	are those whose work operations are or may be in an area where energy control procedures are used. These people must receive instruction regarding the energy control procedures and the prohibition against removing lockout or tagout devices and/or attempting to restart, reenergized or operate the machinery.

Retraining	shall be provided for all authorized and affected employees whenever there is a change in their job assignments, a change in machines, equipment or processes that present a new hazard, or when there is a change in the energy control procedures.

Additional retraining shall also be conducted whenever a periodic inspection reveals, or whenever the employer has reason to believe that there are deviations from or inadequacies in the employee's knowledge or use of the energy control procedures. The retraining shall reestablish employee proficiency and introduce new or revised control methods and procedures, as necessary.

Tagout	The placement of a tagout device on an energy isolating device, in accordance with an established procedure, to indicate that the energy isolating device and the equipment being controlled may not be operated until the tagout device is removed.

Tagout device	A prominent warning device, such as a tag and a means of attachment, which can be securely fastened to an energy isolating device in accordance with an established procedure, to indicate that the energy isolating device and the equipment being controlled may not be operated until the tagout device is removed.

#### References

- 29 CFR 1910.147 The Control of Hazardous Energy (Lockout/Tagout)

- 29 CFR 1910.147(c)(7) Training and Communications

  - Master Lock OSHA Lockout Tagout Best Practices https://www.masterlock.com/service-and-support/videos/safety-solutions/tagout

| Document No.: BioWork-2060 | Revision No.: 01 | Effective Date: 20 Jul 2020 | Page 4 of 13 |
| --- | --- | --- | --- |
| Lockout Tagout Procedure & Demonstration |  |  |  |

#### Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Grainger Master Lock Tool Box : (1) Red/Black Plastic Lockable Carrying Case (Accepts up to 1-1/4" Shackle), (1) 2CJL3 Circuit Breaker Padlock, (1) 1U177 1" Dia. Steel Jaw Hasp, (1) 2CJJ2 Red Labeled Snap-On Hasp, (3) 1TDB2 Standard Toggle Universal Circuit Breaker Lockout, (2) 1TDB1 Tall and Wide Toggle Universal Circuit Breaker Lockout, (1) 2CJK9 Wall Switch Cover, (12) Tags, (3) Red Thermoplastic Keyed-Alike Padlocks | 1 | Toolbox Kit |
| If Kit not available: |  |  |
| Locks | 2 | Each |
| Tags | 2 | Each |
| Multiple Lock Hasps | NS | NS |
| Lockout/Tagout Form (See attachment 9.1) | 1 | Each |
| Emergency Lockout/Tagout Device Removal Form (See attachment 9.1) | 1 | Each |

- Safety

Wear lab PPE including lab coat, gloves, and safety glasses when working in the lab. Follow all identified engineering controls, administrative controls, written procedures and personal protective equipment requirements.

- Utilize all energy isolation devices that have been designed into the equipment being serviced. (Engineering Controls)

- Follow all written procedures, including the use of any necessary lockout or tagout devices. (Administrative Controls)

- Wear all required and specified personal protective equipment which may include but not be limited to safety glasses, safety shoes, gloves, and protective clothing. (Personal Protective Equipment).

| Document No.: BioWork-2060 | Revision No.: 01 | Effective Date: 20 Jul 2020 | Page 5 of 13 |
| --- | --- | --- | --- |
| Lockout Tagout Procedure & Demonstration |  |  |  |

#### Procedure

NOTE: UNDER NO CIRCUMSTANCES SHOULD AUTOMATION BE USED AS A MEANS OF ISOLATION INCLUDING PUTTING PHASES/PROGRAMS IN HOLD.

  - Perform hazard assessment and identify energy sources to be isolated.

    - Where available, use the Hazardous Energy Control Plan (HECP) to identify the types and amounts of energy that power the equipment, and to help understanding the hazards and means of controlling that energy.

    - Notify ALL affected employees and system owner that a LOTO procedure is being used and why it is being locked out by issuing the LOTO form (See attachment 9.1).

    - Authorized Employee complete LOTO FORM Part A and Part B.

    - Shutdown equipment using operating controls.

    - Follow applicable HECP for the equipment to avoid endangering anyone during the shutdown

  - Isolate the equipment/ utility system

    - The Authorized Employee must

      - Notify the System Owner (or designee) prior to shutting down any equipment / utility systems.

    - System Owner (or designee) is to verify:

      - That operations are not being performed on the equipment before shutdown.

      - That all equipment that contained product has been emptied before shutdown.

      - That infectious equipment has been cleaned and sanitized before work is initiated.

      - That all Operators in the area are notified of the planned shutdown.

      - That all identified Lock Out / Tag out locations by the Authorized Employee, are adequate for the task(s) planned.

    - Authorized Employee must then verify:

      - That all energy isolating devices (valves, electrical disconnects, etc.) that are required to control the energy to the equipment/ system must be physically located and operated (closed or shut off) in such a manner as to isolate the equipment/ system from the energy source.

  - Install Locout/Tagout Equipment

    - The points of control must be secured via an isolation device (i.e. locks) so that unauthorized persons are prevented from re-energizing the machine, process or system.

    - Install completed tags on each isolation device with the name of the authorized person, date and contact phone number.

| Document No.: BioWork-2060 | Revision No.: 01 | Effective Date: 20 Jul 2020 | Page 6 of 13 |
| --- | --- | --- | --- |
| Lockout Tagout Procedure & Demonstration |  |  |  |

    - Where a tag cannot be affixed directly to the isolating device, the tag shall be located as close as possible to the device, in a position that will be immediately obvious to anyone attempting to operate the equipment/ utility.

  - Release any stored energy – Potential, additional precautions may include the following:

    - For infectious systems, verify with the system owner that the system has been inactivated and no live virus is present.

    - Inspect the system to make sure all parts have stopped moving.

    - Install ground wires.

    - Release the tension on springs, or block the movement of spring-driven parts.

    - Ensure lines are properly drained and free of liquids.

    - Take any additional precautions to completely isolate the equipment such as the following:

      - Block or brace parts that could fall because of gravity.

      - Block parts in hydraulic and pneumatic systems that could move from loss of pressure.

      - If a line must be blocked where there is no valve, use a blank flange or a cap on a tri-clamp.

      - Dissipate extreme cold or heat wearing proper protective clothing.

      - Release any remaining pressure or steam and confirm said release.

      - If stored energy can re-accumulate, monitor it to ensure it stays below hazardous levels.

  - Verify that the equipment has been properly isolated.

    - For electrical Lockout/Tagout, verify that the main disconnect switch or circuit breaker cannot be moved to the on position.

    - Test the equipment by use of the appropriate test equipment (e.g. voltmeter) and or visual inspection.

    - Press all start buttons and other activating controls on the equipment itself.

    - Shut off all machine controls when the testing is finished.

    - Make certain all dangerous areas are clear of personnel.

    - Monitor, but do not rely upon gauges as the sole means of determining energy presence – re-isolate and check for current where needed!

    - Verify that the equipment cannot be turned to the “on” position. Press all “start” buttons and activating controls on the equipment.

    - Return all controls to the “off” position after certain that Lock Out is effective.

    - Authorized Employee and system owner (or designee) complete Part C of LOTO Form

  - Complete the necessary repairs, interventions, work, manipulations, modifications, etc.

    - Each employee working on the equipment must have his/her own lock and tag placed on the lockout device or on the lock box.

    - Each employee will remove his/her lock and/or tag from the lockout device or lockbox upon completion of his/her portion of the work.

| Document No.: BioWork-2060 | Revision No.: 01 | Effective Date: 20 Jul 2020 | Page 7 of 13 |
| --- | --- | --- | --- |
| Lockout Tagout Procedure & Demonstration |  |  |  |

    - If work will run over multiple shifts, locks can be removed as part of a shift handover. However, the Authorized Employee must leave his/her lock(s) on throughout the entire duration of the work activity, unless they perform an authorization transfer in section A of the LOTO form.

  - Restore the equipment/ utility systems to normal operations.

    - The Authorized Employee must complete Part D on the LOTO Form (FORM

    - Verify no one is exposed to the equipment or machine to be started or operated.

    - Verify all tools have been removed from the machine or equipment.

    - Verify all guards have been reinstalled if removed.

    - Verify there are no exposed electrical wires.

    - Remove all LOTO isolation devices.

    - If an employee is not available to remove his/her lock, see Section 8 for emergency lock removal procedures.

    - Notify system owner (or designee) that the equipment/ utility system is going to be restored to full operation.

    - Re-energize the equipment/ utilities to restore normal operation.

    - Complete Part E of the LOTO Form

    - Authorized Employee and System Owner (or desginee) must sign Part E of the LOTO Form, verifying that all energy sources have been re-energized and authorizing the equipment to be returned to service.

  - Additional Safeguards with Lockout / Tagout

    - Group Lockout/Tagout

      - When multiple personnel are working on the same piece of equipment or utility system, a group LOTO should be utilized.

      - A lead authorized employee must be appointed to coordinate all activities amongst the group.

      - The lead authorized employee will follow this procedure and ensure that all personnel working on the equipment comply with this procedure.

      - The lead Authorized Employee will be the individual to apply the group LOTO to the equipment/system and first individual to apply their lock to the lock box. The lead Authorized Employee is also the last person to remove LOTO devices.

    - Transfer of Lead Authorized Employee Ownership

      - In cases where the lead Authorized Employee cannot be present when the process is complete (off-shift/PTO), then, Lead Ownership may be transferred to the following shift Supervisor or authorized designee capable of being present at the end of the process

      - An “ownership transfer” must be carefully planned requiring the transferred lead to verify/retest the system (Repeat section 8.6) before assuming ownership

    - Emergency Lockout / Tagout Device Removal

| Document No.: BioWork-2060 | Revision No.: 01 | Effective Date: 20 Jul 2020 | Page 8 of 13 |
| --- | --- | --- | --- |
| Lockout Tagout Procedure & Demonstration |  |  |  |

      - If the Authorized Employee who applied the lockout/ tagout device is not available to remove it, the device may be removed under the direction of the Head of Maintenance or Designee.

      - Complete the Emergency LOTO Device Removal Form (Attachment 9.2).

      - The Head of Maintenance or designee must verify the authorized employee who applied the isolation device is not in the facility.

      - The Head of Health, Safety and Environment (or a designee) must pre- approve the removal of the lock.

NOTE: ANY EMPLOYEE WHO REMOVES A LOCKOUT DEVICE OR WARNING TAG WITHOUT AUTHORIZATION, WILL BE SUBJECT TO DISCIPLINARY ACTION UP TO AND INCLUDING TERMINATION

#### Figures and Diagrams

Diagram 1: (If needed, any diagram or drawing)

#### Notes/Observations

- Attachments

  - Lockout/Tagout Form (sample attached)

  - Emergency Lockout/Tagout Device Removal Form (sample attached)

  - Photo of Locks purchased for training exercise

  - Photo of Tags purchased for the training exercise

  - Photo of example of multiple lock hasp

| Document No.: BioWork-2060 | Revision No.: 01 | Effective Date: 20 Jul 2020 | Page 9 of 13 |
| --- | --- | --- | --- |
| Lockout Tagout Procedure & Demonstration |  |  |  |

Sample Lockout / Tagout Form

| Lockout / Tagout Form |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| PART A: Must be completed by the Authorized Employee. |  |  |  |  |  |  |  |
| Work Order #: |  |  | Equipment #: |  |  |  |  |
| Date: |  |  | Equipment Description: |  |  |  |  |
| Authorized Employee (Lead): |  |  | Transferred Lead Authorized Employee: |  |  |  |  |
| Work Description: |  |  |  |  |  |  |  |
| PART B: Identify the energy sources and method of isolation. Issuer must check items are locked and tagged including valve numbers. |  |  |  |  |  |  | PART D: |
| Energy Source | Valve # |  |  | Isolated (Y/N) |  |  | Re -Energized (Y/N) |
| PART C: Authorization to proceed with work, manipulations, etc. – verification that all hazards have been identified and isolated. |  |  |  |  |  |  |  |
| Equipment containing product emptied by System Owner. Yes No NA SO |  | All equipment operations shutdown. Yes No NA SO |  |  | Infectious equipment cleaned and sanitized. Yes No NA SO |  |  |
| All area Operators have been notified of planned shutdown. Yes No NA SO |  | LOTO locations identified are adequate protection for work. Yes No NA SO |  |  | All Energy isolating devices have been closed or shutoff. Yes No NA AE |  |  |
| Authorized Employee: |  |  |  |  |  |  |  |
| System Owner: |  |  |  |  |  |  |  |
| PART E: Authorization to return equipment/ utility to service – verification that all energy sources have been re-energized. |  |  |  |  |  |  |  |
| Verify equipment and area is clear of tools, workers, materials. Yes NA |  | Reposition any safety devices, guards, interlocks Yes NA |  |  | Warned workers to stay clear, verify area is clear of workers. Yes NA |  |  |
| Considered effects of re-energizing on all downstream systems. Yes NA |  | Re-energized System Yes |  |  | Will monitor performance for 24 hrs. prior to product addition. Yes NA |  |  |
| Authorized Employee: |  |  |  |  |  |  |  |

System Owner:

| Document No.: BioWork-2060 | Revision No.: 01 | Effective Date: 20 Jul 2020 | Page 10 of 13 |
| --- | --- | --- | --- |
| Lockout Tagout Procedure & Demonstration |  |  |  |

Sample Emergency LOTO Device Removal Form

In accordance with the Company Hazardous Energy Control Progarm, only the person who applied the LOTO device can remove that LOTO device except as provided in this form. In cases where the Authorized Employee who applied the LOTO device is not available to remove it, the device may only be removed under the direction of the Head of Maintenance (or designee) and preapproval of the Head of Health Safety and Envrionment (or designee). This form should be used only as a last resort to remove a lockout device.

Step 1 – Verify that the employee who applied the lock and tag is not present at the facility. Name of Employee whose lock is to be removed:

| Is the employee in the facility? | Yes / No | Date: Time: |
| --- | --- | --- |
| Is the employee currently employed? | Yes / No | Date: Time: |

Step 2 – Make all reasonable efforts to contact the employee who applied the LOTO device to inform him/her that the lock and/or tag will be removed.

| Has employee contact been attempted? | Yes / No | Date: Time: |
| --- | --- | --- |
| Was contact successful? | Yes / No | Date: Time: |

Step 3 – The Head of Maintenance must verify that the equipment that was locked out is safe to return to service, all personnel involved in the lock out are clear, and no damage will occur when the equipment is restarted.

Has the system been walked down?		Yes / No Date:	Time:	 Is the system safe to return to service?	Yes / No	Date:	Time:

Step 4 - Signature of Head of Maintenance completing steps 1, 2 and 3:

Step 5 - The Site Head of Health, Safety and Environmental (HSE) (or desginee) must authorize the removal of a lock after completion of steps 1, 2, 3 and 4.

Signature of Site HSE (or designee) :	Date:

Step 6 – Record the day and time the lock was removed.

Date and time that the lock was removed: 	/	/	Time:

Step 7 – The Head of Maintenance must ensure that the person who applied the lock and/or tag is notified of the removal before he/she resumes work at the facility.

Step 8 - Date employee notified of lock removal:

Employee:		Date:

| Document No.: BioWork-2060 | Revision No.: 01 | Effective Date: 20 Jul 2020 | Page 11 of 13 |
| --- | --- | --- | --- |
| Lockout Tagout Procedure & Demonstration |  |  |  |

Lockout Tagout Devices

| Model No. 7C5RED Circuit Breaker Compact Padlock |  | Allow proper lockout of all types of circuit breaker lockout devices. Device fits the tight space constraints within circuit breaker shutoff boxes and the 5in (127mm) flexible steel cable allows the compact lock body to be positioned so most panel doors can be closed. |
| --- | --- | --- |
| Model No. 491B Grip Tight™ Circuit Breaker Lockout, Tall and Wide Toggles |  | Lockout Device fits wide or tall breaker toggles typically found on hi-voltage/hi-amperage breakers. The innovative patented Grip Tight™ design effectively locks out circuit breakers with a simple thumb turn and then close clamping handle for a tight grip on the toggle. Device accepts all American Lock® and Master Lock® safety padlocks and most lockout hasps for full lockout. |
| Model No. 493B Grip Tight™ Circuit Breaker Lockout, Standard Single and Double Toggles |  | Lockout Device fits standard height and tie-bar toggles typically found on 120 and 240 volt breakers. The innovative patented Grip Tight™ design effectively locks out circuit breakers with a simple thumb turn and then close clamping handle for a tight grip on the toggle. Device accepts all American Lock® and Master Lock® safety padlocks and most lockout hasps for full lockout. |

| Document No.: BioWork-2060 | Revision No.: 01 | Effective Date: 20 Jul 2020 | Page 12 of 13 |
| --- | --- | --- | --- |
| Lockout Tagout Procedure & Demonstration |  |  |  |

| Model No. 420 Steel Lockout Hasp, 1in (25mm) Jaw Clearance |  | Steel Safety Lockout Hasp features a 1in (25mm) inside jaw diameter and can hold up to six padlocks. Ideal for lockout by multiple workers at each lockout point, the hasp keeps equipment inoperative while repairs or adjustments are made. Control cannot be turned on until last worker’s padlock is removed from hasp. |
| --- | --- | --- |
| Model No. 496B Universal Wall Switch Cover |  | Universal Wall Switch Cover locks out both toggle and Decora™ switches. The patented design ensures switch stays locked on or off and hinges open for unobstructed normal operations of switch. |
| Model No. 427 Labeled Snap-on Lockout Hasp, Red, 1-3/4in x 2-1/8in (44.5mm x 54mm) Jaw Clearance |  | Lockout Hasp features a 1- 3/4in x 2-1/8in (45mm x 54mm) diameter inside jaw clearance and is designed to fit more lockout points than alternative rivet-hinge labeled lockout hasps. The anodized aluminum and stainless steel construction provides combined corrosion resistance and strength. A spring-loaded locking arm offers an easy-to-apply "snap- on" advantage. Includes high- visibility, write-on safety labels |

| Document No.: BioWork-2060 | Revision No.: 01 | Effective Date: 20 Jul 2020 | Page 13 of 13 |
| --- | --- | --- | --- |
| Lockout Tagout Procedure & Demonstration |  |  |  |

| Model No. 497A Do Not Operate Safety Tag, English, Laminated |  | Heavy Duty Safety Tags feature durable polyester laminate to resist water, grease and extreme temperatures. The heavy duty construction withstands over 50 lbs. (23kg) of pull force. The 7/16in (11mm) diameter brass grommet accepts all safety padlocks. |
| --- | --- | --- |
| Model No. 410KARED Red Zenex™ Thermoplastic Safety Padlock, 1-1/2in (38mm) Wide with 1-1/2in (38mm) Tall Shackle |  | Designed exclusively for Lockout/Tagout applications, the durable, lightweight, non- conductive lock body is easy to carry and padlock features high security, reserved-for-safety cylinder with key retaining, to ensure padlock is not left unlocked. |

| Document No.: BioWork-3030 | Revision No.: 05 | Effective Date: 22 Aug 2019 | Page 1 of 9 |
| --- | --- | --- | --- |
| Pipette and Micropipette |  |  |  |

#### Purpose

This procedure will define the proper use of pipettes and micropipettes in the laboratory.

#### Scope

This procedure applies to the equipment and operation of various pipetting and micropipetting equipment to provide technicians with an understanding of the proper care and use of the appropriate equipment for small-volume liquid measurement.

This procedure applies to lab or processing activities that require precise measurement of small volumes of liquids.

Note: This is the note section.

#### Responsibilities

Technicians will only operate pipettes, micropipette, and their associated equipment under the direction of an instructor. It is the responsibility of students undergoing gowning training to comply with this procedure.

#### Definitions

Adjustment: Altering the micropipette so that the dispensed volume is the desired amount.

Air Displacement Pipette: Micropipettes meant for general use with aqueous solutions. When using this type of pipette, a certain volume of air remains between the piston and the liquid

Aspirate: A verb used to describe the action of drawing sample solutions into the pipette tube or tip

AU: Absorbance Units. A numerical value for the amount of light absorbed by the sample solution when using a spectrophotometer

Blow-out: To empty the tip completely

BPR: Batch Production Record

Dispense: A verb used to describe the action of expelling the sample solution from the pipette tube or tip into an appropriate receiving container

Diluent / Dilution Buffer: A substance that dilutes the strength of a solution or mixture

NLT:Not Less Than

μl: microliter. A volume designation used for the calibration and use of micropipettes. 1000.0 μl is equal

to 1.0 mL

| Document No.: BioWork-3030 | Revision No.: 05 | Effective Date: 22 Aug 2019 | Page 2 of 9 |
| --- | --- | --- | --- |
| Pipette and Micropipette |  |  |  |

Micropipette: A mechanical device used to consistently measure identical volumes according to the setting on the device, typically very small volumes.

Pipette: A tubular device used to accurately measure volumes of liquids within the stated range of the pipette

Pipette Bulb: Devices used to facilitate aspirating and dispensing of sample materials within a pipette.

Never pipette by mouth.

Positive Displacement (PD) Pipette: Used for high viscosity and volatile liquids. In this type of pipette, the liquid is in direct contact with the piston

SOP: Standard Operating Procedure

Tip:	A removable device placed onto the end of a micropipette to contain the sample solution being aspirated/dispensed. Except when using PD pipettes, the tip should have an air gap between the sample and the micropipette to avoid fouling/contaminating the micropipette.

#### References

- Gutzman, K., Pipetting: A Practical Guide, The American Biology Teacher, Vol. 63, NO. 2, February 2001

#### Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Lab Water | 100 | mL |
| 250 mL beaker (for diluent) | 1 | Each |
| 1L beaker (for diluent) | 1 | Each |
| Micro-centrifuge Tubes (1.5 mL capacity) | 3 | Each |
| 10 mL Volumetric Pipette and pipette bulb (pipette aid) | 1 | Each |
| Micropipettes: P20, P200, and P100 | 1 | Each |
| Micropipette tips (2 sizes) | 3 | Each |
| Waste Container | 1 | Each |

| Document No.: BioWork-3030 | Revision No.: 05 | Effective Date: 22 Aug 2019 | Page 3 of 9 |
| --- | --- | --- | --- |
| Pipette and Micropipette |  |  |  |

- Safety

Wear lab PPE including lab coat, gloves, and safety glasses. Never mouth pipette; always use a pipette bulb.

#### Procedure General Notes

  - Principles of Operation

- Pipettes and micropipettes are used to aspirate and dispense accurate volumes of liquid samples.

- Technicians should exercise caution when aspirating solutions to avoid wetting the interior of the pipette bulb or micropipette.

- Identify the correct size pipette to be used based on the volume to be measured. Choose a pipette with a capacity larger than, but as close to the required sample volume as possible.

- Attach the pipette to a pipette bulb to facilitate aspiration and dispensing in a safe, controlled manner.

  - Choosing the Correct Micropipette

    - Each micropipette has a range of volumes for which it is accurate and reliable. This range is stamped on the pipette below the display window and should never be exceeded.

    - Selection of the correct pipette is crucial to accurately pipetting the desired volume. The range of variable volumes pipettes you may encounter include:

- P2 (0.5-2.0µl)

- P10 (1-10 μl)

- P20 (2-20 μl)

- P200 (20- 200 μl)

- P1000 (200-1000 μl)

- P5000 (1000-5000 μl)

    - Each type of micropipette requires a tip that fits the end of the pipette correctly.

    - Determine the volume you need to pipet and then select the correct pipette for the task.

    - Set the micropipette to the required volume by rotating the plunger button on the end until the correct numbers appear in the display.

| Document No.: BioWork-3030 | Revision No.: 05 | Effective Date: 22 Aug 2019 | Page 4 of 9 |
| --- | --- | --- | --- |
| Pipette and Micropipette |  |  |  |

    - Select and fit the appropriate-size tip for the micropipette selected.

  - Procedure for pipetting 5 mL using a pipette bulb and 10 mL pipette

    - Insert a 10 mL pipette into the pipette bulb.

    - Insert the tip of the pipette into the sample solution a few millimeters below the surface, but deep enough so that you don’t draw up air as the solution level in the sample container decreases.

    - While viewing the graduations at eye-level and holding the pipette vertically, aspirate the required volume of solution into the pipette.

    - Once the required volume has been drawn into the pipette, withdraw the pipette from the sample solution while dragging the tip up the wall of the container to remove excess sample material from the outside of the pipette.

    - Place the tip of the pipette against the interior wall of the receiving container and dispense the solution slowly.

Caution: forceful dispensing of samples can cause the formation of aerosols, a potential safety hazard and source of contamination

Note: the pipette may be reused immediately for the same sample solution. If a different solution must be sampled or the pipette will not be reused immediately, place the used pipette into the designated waste container and acquire a new pipette for the next operation.

  - Use a micropipette to deliver 15 uL

    - Obtain a P20 micropipette.

    - Set the volume to 15 uL.

    - Attach a tip.

    - Transfer 15 uL of solution (lab water) into a microcentrifuge tube.

      - Label the micro-centrifuge tube as “Tube 1.” and weigh the empty tube. Record the empty tube weight in the table below.

      - Depress the pipette to the first stop.

| Document No.: BioWork-3030 | Revision No.: 05 | Effective Date: 22 Aug 2019 | Page 5 of 9 |
| --- | --- | --- | --- |
| Pipette and Micropipette |  |  |  |

      - While still holding the plunger button down, hold the pipette vertically and immerse the tip into the liquid and hold it at a constant depth below the surface of the liquid.

Note: remember to account for the volume to be withdrawn when inserting the pipette tip into the sample solution to prevent aspirating air. Aspiration of air will negatively affect the accuracy of the sample volume

      - Slowly release the plunger button so that the pipette will draw the liquid into the tip.

Note: Never allow the plunger to “snap” back as it can affect the accuracy of the micropipette. As the depth of liquid lowers, so must the tip of the pipette.

      - Once the plunger button has been completely released, hold the pipette in the liquid for 1 additional second to complete aspiration.

      - Remove the pipette and tip from the solution while dragging the tip against the interior wall of the container to remove excess sample material from the outside of the tip. Continue to hold the pipette vertically.

      - To deliver the sample, place the tip against the inside wall of the microcentrifuge tube at a slight angle. Press the plunger button slowly and smoothly to the first stop.

      - Wait at least one second as it dispenses and then press the plunger button further to the second stop to blow out any residual liquid from the tip to deliver (TD) the desired volume.

      - Keep the plunger button pushed down as you withdraw your tip from the micro-centrifuge tube so as not to re-aspirate any of the transferred sample.

      - Release the plunger button slowly and smoothly.

    - Repeat steps 8.4.4.2 – 8.4.4.10 nine more times for a total of 10 transfers to same tube.

    - Remove the pipette tip

      - The tip may be ejected from the micropipette barrel by pressing the tip ejector button while holding the pipette tip over a waste container or by gently twisting and pulling the tip from the micropipette barrel. Place the used tip in an approved waste container.

| Document No.: BioWork-3030 | Revision No.: 05 | Effective Date: 22 Aug 2019 | Page 6 of 9 |
| --- | --- | --- | --- |
| Pipette and Micropipette |  |  |  |

    - Close the filled micro-centrifuge tube and obtain the filled weight. Record the filled weight in the table below.

  - Use a micropipette to deliver 150 uL

    - Obtain a P200 micropipette.

    - Set the volume to 150 uL.

    - Attach a tip.

    - Transfer 150 uL of solution (lab water) into a microcentrifuge tube.

      - Label the micro-centrifuge tube as “Tube 2.” and weigh the empty tube. Record the empty tube weight in the table below.

      - Depress the pipette to the first stop.

      - While still holding the plunger button down, hold the pipette vertically and immerse the tip into the liquid and hold it at a constant depth below the surface of the liquid.

      - Slowly release the plunger button so that the pipette will draw the liquid into the tip.

      - Once the plunger button has been completely released, hold the pipette in the liquid for 1 additional second to complete aspiration.

      - Remove the pipette and tip from the solution while dragging the tip against the interior wall of the container to remove excess sample material from the outside of the tip. Continue to hold the pipette vertically.

      - To deliver the sample, place the tip against the inside wall of the microcentrifuge tube at a slight angle. Press the plunger button slowly and smoothly to the first stop.

      - Wait at least one second as it dispenses and then press the plunger button further to the second stop to blow out any residual liquid from the tip to deliver (TD) the desired volume.

      - Keep the plunger button pushed down as you withdraw your tip from the microcentrifuge tube so as not to re-aspirate any of the transferred sample.

| Document No.: BioWork-3030 | Revision No.: 05 | Effective Date: 22 Aug 2019 | Page 7 of 9 |
| --- | --- | --- | --- |
| Pipette and Micropipette |  |  |  |

      - Release the plunger button slowly and smoothly.

      - Repeat steps 8.5.4.2 – 8.5.4.10 nine more times for a total of 10 transfers to same tube (totaling 1,500 uL).

    - Remove the pipette tip

      - The tip may be ejected from the micropipette barrel by pressing the tip ejector button while holding the pipette tip over a waste container or by gently twisting and pulling the tip from the micropipette barrel. Place the used tip in an approved waste container.

    - Close the filled micro-centrifuge tube and obtain the filled weight. Record the filled weight in the table below.

  - Use a micropipette to deliver 500 uL

    - Obtain a P1000 micropipette.

    - Set the volume to 500 uL.

    - Attach a tip.

    - Transfer 500 uL of solution (lab water) into a micro-centrifuge tube.

      - Label the micro-centrifuge tube as “Tube 3.” and weigh the empty tube. Record the empty tube weight in the table below.

      - Depress the pipette to the first stop.

      - While still holding the plunger button down, hold the pipette vertically and immerse the tip into the liquid and hold it at a constant depth below the surface of the liquid.

      - Slowly release the plunger button so that the pipette will draw the liquid into the tip.

      - Once the plunger button has been completely released, hold the pipette in the liquid for 1 additional second to complete aspiration.

      - Remove the pipette and tip from the solution while dragging the tip against the interior wall of the container to remove excess sample material from the outside of the tip. Continue to hold the pipette vertically.

| Document No.: BioWork-3030 | Revision No.: 05 | Effective Date: 22 Aug 2019 | Page 8 of 9 |
| --- | --- | --- | --- |
| Pipette and Micropipette |  |  |  |

      - To deliver the sample, place the tip against the inside wall of the micro- centrifuge tube at a slight angle. Press the plunger button slowly and smoothly to the first stop.

      - Wait at least one second as it dispenses and then press the plunger button further to the second stop to blow out any residual liquid from the tip to deliver (TD) the desired volume.

      - Keep the plunger button pushed down as you withdraw your tip from the micro-centrifuge tube so as not to re-aspirate any of the transferred sample.

      - Release the plunger button slowly and smoothly.

      - Repeat steps 8.6.4.2 – 8.6.4.10 two more times for a total of 3 transfers (totaling 1,500 uL).

    - Remove the pipette tip

      - The tip may be ejected from the micropipette barrel by pressing the tip ejector button while holding the pipette tip over a waste container or by gently twisting and pulling the tip from the micropipette barrel. Place the used tip in an approved waste container.

    - Close the filled micro-centrifuge tube and obtain the filled weight. Record the filled weight in the table below.

  - Visually compare the final volumes of tubes 2 and 3. Determine (visually) if they are close in volume. Compare the calculated net weights. Document your findings in the Notes/Observations section below, or in your lab notebook.

N/A

#### Figures and Diagrams

- Notes/Observations

| Document No.: BioWork-3030 | Revision No.: 05 | Effective Date: 22 Aug 2019 | Page 9 of 9 |
| --- | --- | --- | --- |
| Pipette and Micropipette |  |  |  |

Fill Weights of micro-centrifuge tubes

|  | TUBE #1 | TUBE #2 | TUBE #3 |
| --- | --- | --- | --- |
| Volume per Transfer | 15µL | 150µL | 500µL |
| Number of Transfers | 10 | 10 | 3 |
| Expected Net Weight (g) | 0.150 g | 1.50 g | 1.50 g |
| Empty Tare Weight (g) |  |  |  |
| Gross (Filled) Weight (g) |  |  |  |
| Actual Net Weight Gross wt. – Tare wt.) g |  |  |  |
| Percent Error |  |  |  |
| Performed By: Date: |  |  |  |
| Verified By: Date: |  |  |  |

Note: Assume 1cc (1mL) of water = 1.0 g

Percent Error = (Difference of expected weight from actual weight) ÷ Expected weight , Changed to percent

#### Example:

Expected Weight = 0.150 g Actual Weight

= 0.148 g

Percent Error  = 0.150 – 0.148 = 0.002

= 0.002 ÷ 0.150 = 0.0133

= 1.33% Error

N/A

#### Attachments

| Document No.: BioWork-3050 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 1 of 7 |
| --- | --- | --- | --- |
| Measuring Weight |  |  |  |

- Purpose

To familiarize with the types of scales/balances, expected working ranges, and to practice measuring weight.

#### Scope

This activity has three parts:

- Part 1: Choosing Your Instrument

- Part 2: Verification of a Balance

- Part 3: Measuring Mass in Commercial Packaging

#### Responsibilities

Technicians will operate equipment under the direction of an instructor. It is the responsibility of students undergoing gowning training to comply with this procedure.

#### Definitions

Instrument error:	The difference between the indicated value (the measurement shown on the instrument) and the true value of the object being measured.

#### References

N/A

#### Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Balances (Various models) | 2 to 3 | Each |
| Nitrile Gloves | 1 | Pair |
| Large Weigh Boats | 3 | Each |
| Small Weigh Boats | 3 | Each |
| Spatulas | 3 | Each |
| Standard Weights | 1 | Set |
| Sample Package Items; Examples: pkg. snack crackers, small box of raisins, pkg. of chewing gum etc. | 3 | Each |

| Document No.: BioWork-3050 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 2 of 7 |
| --- | --- | --- | --- |
| Measuring Weight |  |  |  |

- Safety

Wear lab PPE including lab coat, gloves, and safety glasses.

#### Procedure

  - Part 1 – Choosing Your Instrument (Lab Balance Inventory)

    - Examine the balances in your lab. Determine the range of each one and record it below. If the balance is not marked with a range, label it.

Note: See Attachment for a table of balances currently used in the BioNetwork.

| Balance |  | Range |
| --- | --- | --- |
| Manufacture | Model |  |
| Performed By: Date: |  |  |
| Verified By: Date: |  |  |

    - Using the table below, select which lab balance (from the list above) that would be appropriate to use for the mass measurements listed in the table. If you do not have a balance that is accurate for a particular mass, write “out of range”.

Note: Taring out a weight does not remove it from the gross weight capacity. The maximum capacity of the balance includes the weight boat plus the material you’re measuring.

| Document No.: BioWork-3050 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 3 of 7 |
| --- | --- | --- | --- |
| Measuring Weight |  |  |  |

| Weight to be Measured | Balance Selected for Measurement (Type and Asset Tag #) |
| --- | --- |
| 0.1 g |  |
| 1 g |  |
| 10 g |  |
| 100 g |  |
| 1 kg |  |
| Performed By: Date: |  |
| Verified By: Date: |  |

  - Part 2 – Verification of a Balance

    - Before using a measurement instrument like a balance, you must be sure that it is giving an appropriate reading by using standard weights. The standard weights you use should match the range (above and below) of your intended measurements, and within the range of the balance. Typically your SOP will provide an allowable deviation.

    - Ensure that the balance is placed on a level, stable surface and is clean.

    - Turn on the balance, if it is not on already. The digital readout should light up.

    - Next, make sure that the reading on the balance is zero. If it is not, use the tare button to adjust the readout to zero.

    - Check the balance by weighing two standard weights - one below and one above the intended measurement.

Note: You must wear cotton or nitrile gloves and /or use the forceps to handle the standard.

    - Record the specified weight of the lower standard in the Balance Verification table below.

| Document No.: BioWork-3050 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 4 of 7 |
| --- | --- | --- | --- |
| Measuring Weight |  |  |  |

    - Record the specified weight of the higher standard in the Balance Verification table below.

    - Ask the instructor what the allowable deviation is for each standard and record it in the Balance Verification table below

    - Place each standard on the appropriate balance and record their indicated (actual) weight from the readout, in the Balance Verification table below.

    - Calculate the deviation by subtracting the actual weight (indicated) from the specified weight and record it in the Balance Verification table below.

    - If the reading on the scale matches the standard weights within the allowable deviation, check pass.

    - If the reading on the scale does not match the standard weights, check fail. (If your balance fails, notify your instructor.)

Balance Verification Table

| Standard Weights |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Standard | Specified Weight (mg or g) | Actual Weight (mg or g) | Allowable Deviation (mg or g) | Actual Deviation (mg or g) | Pass/Fail |
| Lower Std |  |  |  |  |  |
| Above Std |  |  |  |  |  |
| Performed By: Date: |  |  |  |  |  |
| Verified By: Date: |  |  |  |  |  |

  - Part 3 – Measuring Mass in Commercial Packaging

    - Ensure that the balance is placed on a level, stable surface and is clean.

    - Turn on the balance, if it is not on already. The digital readout should light up.

    - Next, make sure that the reading on the balance is zero. If it is not, use the tare button to adjust the readout to zero.

| Document No.: BioWork-3050 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 5 of 7 |
| --- | --- | --- | --- |
| Measuring Weight |  |  |  |

Briefly examine the samples that you are going to weigh and, for each sample, record on the table in the Notes/Observations Section.

- Name of sample

- Type of balance you will use, its range, and asset tag number

- Manufacturer’s weight printed on the package

    - Use the following procedure to measure each sample and complete the chart.

      - Tare the empty weighing container that will hold your sample. (Taring subtracts the weight of the empty container from the total weight of the container plus sample.)

      - Empty the sample from the package (box, bottle, etc.) into the weighing container. Make sure you get the entire sample out of the package.

      - Weigh and record the actual weight on the form.

Calculations

      - Compare the Net weight printed on the package to the weight on the balance.

      - If there is a difference (deviation), make sure you removed all residue from the package. If there is still a difference, calculate the percent difference. To calculate the percent difference:

        - Subtract the smaller weight from the larger weight to get the difference.

        - Divide the difference by the larger weight to get the percent difference.

        - For example, 10 grams is printed on the package but the balance weighs the sample at 9.5 grams so 0.5 grams is the difference and 0.5 divided by 10 = 0.05, which is 5%.

| Document No.: BioWork-3050 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 6 of 7 |
| --- | --- | --- | --- |
| Measuring Weight |  |  |  |

#### Figures and Diagrams

N/A

#### Notes/Observations

Measuring Mass in Commercial Packaging

| Name of Sample | Type of Balance (Asset #, Range) | Weight on the Package (Net Weight) | Actual Weight of Pkg. Contents | Deviation in weight | Percent Difference |
| --- | --- | --- | --- | --- | --- |
| Performed By: Date: |  |  |  |  |  |
| Verified By: Date: |  |  |  |  |  |

  - Explain possible reasons for any deviations noted:

| Document No.: BioWork-3050 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 7 of 7 |
| --- | --- | --- | --- |
| Measuring Weight |  |  |  |

#### Attachments

Table 1: Examples of BioNetwork Scales/Balances

| Balance |  |  | Range |  |
| --- | --- | --- | --- | --- |
| Manufacture | Model | Type | Max Weight | Resolution (d) |
| Fisher | AMF 204 | Analytical | 210 g | 0.0001 g |
| Mettler Toledo | PL 202-S | Top Load | 210 g | 0.01 g |
| Mettler Toledo | ME 2000 2E | Top Load | 2.2 Kg | 0.01 g |
| Fisher | ALF 1202 | Top Load | 1200 g | 0.01 g |
| Ohaus | AV 812 | Top Load | 810 g | 0.01 g |
| Ohaus | VP 2102 CN | Top Load | 2100 g | 0.01 g |
| Ohaus | CD-33 | Floor | 10 Kg | 10 g (0.01 Kg) |

| Document No.: BioWork-3060 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 1 of 9 |
| --- | --- | --- | --- |
| Calibration and Operation of the Seven Excellence pH/conductivity Benchtop Meter |  |  |  |

#### Purpose

This document will define how to operate and calibrate the Seven Excellence pH/conductivity benchtop meter.

#### Scope

This procedure applies to all Mettler Toledo, Seven Excellence benchtop pH/conductivity meters.

The Seven Excellence pH/conductivity meter for benchtop measurements will be calibrated at the beginning of each program and recorded in the appropriate activity batch record or log book.

#### Responsibilities

Personnel will operate and calibrate the Seven Excellence benchtop pH/conductivity meter under the supervision of an instructor. It is the responsibility of students undergoing training to comply with this procedure.

#### Definitions

KCl:	Potassium Chloride

NIST: National Institute of Standards and Technology

NLT:	Not Less Than

NMT:	Not More Than

WFI:	Water for Injection

mS/cm:	millisiemens per centimeter

µS/cm:	microsiemens per centimeter

#### References

- Mettler Toledo Seven Excellence

https://www.mt.com/us/en/home/products/Laboratory_Analytics_Browse/pH- meter/benchtop-pH- meter/sevenexcellence.html?cmp=sea_05010127&SE=GOOGLE&Campaign=MT_ANA- pH_EN_US&Adgroup=PH+Meter+-+Benchtop+-

+SevenExcellence&bookedkeyword=%2Bseven%20%2Bexcellence%20%2Bph%20%2Bme ter&matchtype=b&adtext=241566963515&placement=&network=g&kclid=_k_EAIaIQobChM IjdTL34aP4gIVEVuGCh2xhw- qEAAYASAAEgLrD_D_BwE_k_&gclid=EAIaIQobChMIjdTL34aP4gIVEVuGCh2xhw- qEAAYASAAEgLrD_D_BwE

| Document No.: BioWork-3060 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 2 of 9 |
| --- | --- | --- | --- |
| Calibration and Operation of the Seven Excellence pH/conductivity Benchtop Meter |  |  |  |

#### Materials/Equipment

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Mettler Toledo Seven Excellence pH/conductivity meter | 1 | Each |
| InLab Routine Pro pH Electrode (Temperature compensated pH electrode) | 1 | Each |
| InLab 731-ISM conductivity electrode (Measures mS/cm) black electrode | 1 | Each |
| InLab 741-ISM conductivity electrode (Measures µS/cm) silver electrode | 1 | Each |
| Printer paper (continuous label) | 1 | Roll |
| Water bath, typically calibrated to maintain 25°C (± 2°C) (if needed). | 1 | Each |
| Calibrated thermometer (Instrumentation calibrated) | 1 | Each |
| pH 7.00 Buffer - Any commercial phosphate buffer solution accurate to ± 0.01 pH units and traceable to NIST | 20 | mL |
| pH 4.00 Buffer - Any commercial Phthalate buffer solution accurate to ± 0.01 pH units and traceable to NIST | 20 | mL |
| pH 6.86 Buffer - Any commercial buffer solution accurate to ± 0.01 pH units and traceable to NIST | 20 | mL |
| Water: Purified water, USP/EP or better. Example: WFI (tempered) | 100 | mL |
| Suitable pH electrode electrolyte solution (KCl based) | 50 | mL |
| NIST traceable 1413 µS/cm solution | 50 | mL |

- Safety

| Document No.: BioWork-3060 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 3 of 9 |
| --- | --- | --- | --- |
| Calibration and Operation of the Seven Excellence pH/conductivity Benchtop Meter |  |  |  |

Use the appropriate personal protective equipment (PPE) and follow Lab safety guidelines for handling acids, bases, and organic solvents. PPE may include wearing and the proper use of each of the following: gloves, apron, facemask, face shield, or goggles.

#### Procedure

General Notes

- Store the pH electrode in a 7.00 pH buffer or suitable pH electrode storage solution when not in use. This pH storage solution must be changed once per shift.

- The conductivity electrodes are stored dry. Do not submerge the electrode in any solution between uses.

- The electrodes must be replaced when they malfunction, fail to calibrate, or if the response time consistently takes longer than one minute. Consider placing a new pH electrode into service whenever the slope falls below 95% or approximately every year (whichever occurs first). Consider replacement of the conductivity probe whenever the conductivity electrode cannot be calibrated

- All electrodes are to be rinsed with WFI and blotted dry prior to submerging into any solution.

  - Replacement of Electrodes (conditioning new electrode)

    - InLab Routine Pro pH Electrode

      - Remove the protective tip cap

      - Clean the exterior by rinsing with WFI

      - Uncover the KCl filling hole by removing the plug. To maintain adequate volume of the KCl the filling solution must be NMT 1/4 inch below the fill hole. The filling hole should be open when in use.

      - Shake down the electrode to remove internal air bubbles. Pre-condition in a 7.00 pH buffer solution for NLT 1 hour.

    - Seven Excellence InLab 731-ISM or InLab 741-ISM conductivity electrode

      - Remove all packaging material from sensor before use.

      - Clean the exterior by rinsing with WFI

  - Two point Calibration method for the pH module

    - The buffers chosen for the calibration should bracket the samples assayed.

    - The aliquots of calibration buffers that are used for the daily calibrations must be changed at the beginning of each shift. Verify that the buffer solutions are not outdated.

    - For pH measurements the 7.00 pH and the 4.00 pH buffers are chosen for the calibration. System verification is performed using a 6.86 pH buffer.

    - From the home screen tap the Calib pH icon.

    - Rinse the pH electrode in use with WFI, blot dry with a lint-free tissue.

| Document No.: BioWork-3060 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 4 of 9 |
| --- | --- | --- | --- |
| Calibration and Operation of the Seven Excellence pH/conductivity Benchtop Meter |  |  |  |

    - Tap the start icon and follow the on screen instructions.

    - Rinse probe in between readings.

NOTE: During the calibration the letter (A) will continue to flash until the meter locks in on the correct measurement.

    - Tap OK and verify that the printer prints out the readings.

    - Rinse the electrode after calibration, blot dry with a lint-free tissue, and place into the 6.86 pH buffer solution. Tap the “Measure pH icon”. Tap the “Start icon”.

      - During the reading the letter (A) will continue to flash until the meter locks in on the correct measurement.

    - Then tap the OK icon.

    - The printer will print a read out. Verify from the printout that all parameters are correct as required by the log book and tape or glue the printed label in the log book. The 6.86 buffer solution readings must be within ± 0.02 pH units. (If the buffer is not within range, re-calibrate the module). Sign your initials across the edge of the label in the log book (attached).

    - Document the calibration in the pH/conductivity meter logbook.

      - If the calibration fails to meet the system suitability requirements, end task and perform required maintenance and document corrective action taken and acceptable re-calibration of the meter in the appropriate sections in the logbook.

    - Rinse the electrode and blot dry then submerge it into the 7.00 storage solution.

  - Performing a pH measurement using pH module

    - Prepare the sample (dilute and adjust temperature) as specified in the applicable procedure.

    - Prepare the pH electrode by removing the electrode from the storage buffer solution. Rinse the electrode with WFI and blot the electrode dry. Do not wipe.

    - Place the pH electrode in the sample. Tap the “Measure pH icon” from the Home screen then tap the “Start” icon. Start the measurement and wait for it to stabilize.

NOTE: During the reading the letter (A) will continue to flash until the meter locks in on the correct measurement.

    - Once the meter has locked in tap the OK icon and the printer will print the final solution measurements. Place this label into the attached Seven Excellence Meter Calibration Logbook.

    - When taking multiple readings while adjusting a buffer up or down it is not necessary to print a label out each time, only on the final adjusted pH is it necessary to print out a label.

      - To take a pH without printing out a label, tap the check box under the pH/mV icon (make sure that only the pH module is selected and not the

| Document No.: BioWork-3060 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 5 of 9 |
| --- | --- | --- | --- |
| Calibration and Operation of the Seven Excellence pH/conductivity Benchtop Meter |  |  |  |

other modules), and then tap the read icon at the bottom right hand side of the display.

      - Once complete, tap the OK icon to return to the Home screen.

    - Remove the pH electrode from the sample. Discard the sample, rinse the electrode with WFI and blot it dry. Do not wipe. Insert the electrode into the 7.00 pH storage buffer solution.

  - Procedure for operating the conductivity modules

NOTE: This is a single point Calibration method for the conductivity modules (the mS/cm and µS/cm conductivity modules are both calibrated the same way and with the same solution). Both modules (if attached) will need to be calibrated independently. The InLab 731-ISM conductivity electrode (Measures mS/cm) and is the black electrode. The InLab 741-ISM conductivity electrode (Measures µS/cm) and is the silver electrode.

    - Calibrating the mS/cm module

      - From the home screen tap the Cal mS/cm icon.

      - Rinse the conductivity electrode in use (the black electrode measures mS/cm) with WFI, blot dry with a lint-free tissue. Immerse the electrode into the 1413 µS/cm solution, deep enough to submerge the gaps at the bottom of the electrode.

      - Tap the Start icon and follow the on screen instructions.

NOTE: During the calibration the letter (A) will continue to flash until the meter locks in on the correct measurement

      - Tap OK and verify that the printer prints out the readings.

    - Document the Calibration in the pH/conductivity meter logbook. Verify from the printout that all parameters are correct as required by the log book and tape or glue the printed label in the log book. Sign your initials across the edge of the label in the log book.

      - If the calibration fails to meet the system suitability requirements, perform required maintenance and document corrective action taken and acceptable re-calibration of the meter in the appropriate sections in the logbook.

    - Prior to calibrating the µS/cm module, take a suitability check for the mS/cm module. This may be printed and attached to activity batch record.

    - If there is not a µS/cm module connected, proceed to section 8.2.2. Otherwise, , rinse the electrode and blot dry. The electrode will be stored suspended and dry.

  - Performing a conductivity measurement using the Conductivity module

NOTE: This section applies to measuring conductivity using mS/cm.

| Document No.: BioWork-3060 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 6 of 9 |
| --- | --- | --- | --- |
| Calibration and Operation of the Seven Excellence pH/conductivity Benchtop Meter |  |  |  |

    - Rinse the conductivity electrode with WFI and blot the electrode dry. Immerse the electrode into the solution to be measured.

    - Tap the Msr mS/cm icon from the Home screen then, tap the Start icon. Start the measurement and wait for it to stabilize.

NOTE: During the reading the letter (A) will continue to flash until the meter locks in on the correct measurement.

    - Once the meter has locked in, tap the OK icon and the printer will print the final solution measurements. Place this label into the attached Seven Excellence Meter Calibration Logbook.

    - When taking multiple readings while adjusting a buffer up or down it is not necessary to print a label out each time, only on the final adjusted conductivity is it necessary to print out a label.

      - To take a conductivity reading without printing out a label, tap the check box under the appropriate conductivity icon (make sure that only the desired conductivity module is selected and not the other modules), and then tap the read icon at the bottom right hand side of the display.

      - Once complete, tap the OK icon to return to the Home screen.

    - Remove the conductivity electrode from the sample. Discard the sample, rinse the electrode with WFI and blot it dry. Do not wipe. Store the conductivity electrodes elevated and dry.

    - Record daily calibration and calibration data in the attached Seven Excellence Meter Calibration Logbook or equivalent.

    - These are substeps

#### Figures and Diagrams

Mettler Toledo Seven Excellence

| This is a sub substepDocument No.: BioWork-3060 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 7 of 9 |
| --- | --- | --- | --- |
| Calibration and Operation of the Seven Excellence pH/conductivity Benchtop Meter |  |  |  |

#### Notes/Observations N/A

| Document No.: BioWork-3060 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 8 of 9 |
| --- | --- | --- | --- |
| Calibration and Operation of the Seven Excellence pH/conductivity Benchtop Meter |  |  |  |

- Attachments

11.1	Seven Excellence Meter Calibration Logbook (Example)

| pH module | pH Buffers Within Expiration | 7.00 pH Buffer Standardized | 4.00 pH Buffer Standardized | Slope (95%- 102%) | 6.86 pH Buffer (6.84 to 6.88) | Initials/date |
| --- | --- | --- | --- | --- | --- | --- |
| [ ] N/A [ ] Yes | [ ] Yes | [ ] Yes | [ ] Yes | [ ] Yes | [ ] Yes |  |
| mS/cm Conductivit y module | Standard solution within expiration: Check date on Sachet | Calibrated with 1413 µS/cm solution | Suitability Check (1.39 to 1.44 mS/cm) | N/A | N/A | Initials/date |
| [ ] N/A [ ] Yes | [ ] Yes | [ ] Yes | [ ] Yes | N/A | N/A |  |
| µS/cm Conductivit y module | Standard solution within expiration: Check date on Sachet | Calibrated with 1413 µS/cm solution | Suitability Check (1385 to 1441 µS/cm) | N/A | N/A | Initials/date |
| [ ] N/A [ ] Yes | [ ] Yes | [ ] Yes | [ ] Yes | N/A | N/A |  |

Note: Refer to SOP for Seven Excellence pH/Conductivity Meter operating instructions.

Note: Buffer expiration is one month from in service date.

| Document No.: BioWork-3080 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 1 of 4 |
| --- | --- | --- | --- |
| Measuring Density |  |  |  |

#### Purpose

To practice measuring and expressing density in metric units.

#### Scope

This activity has three parts:

- Measuring the Density of Common Liquids

- Measuring the Density of Common Solids

- Measuring Volume (displacement) to Determine Density

#### Responsibilities

It is the responsibility of students undergoing gowning training to comply with this procedure. Technicians will operate equipment under the direction of an instructor.

- Definitions N/A

- References N/A

- Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| 50-mL graduated cylinders | 4 | Each |
| 100-mL or 250ml graduated cylinders | 2 | Each |
| Top Load Balance (scale) or equivalent | 1 | Each |
| Large Weigh Boats | 3 | Each |
| Spatula | 1 | Each |
| Disposable Transfer Pipettes (10 ml) | 3 | Each |

| Document No.: BioWork-3080 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 2 of 4 |
| --- | --- | --- | --- |
| Measuring Density |  |  |  |

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Liquid Samples (at least 2 of different densities), examples: - 16% w/w Sucrose - 95% Ethanol | 2 | Each |
| Solid Samples (at least 2 of different densities), examples: - Aquarium Gravel - Lead sinkers (split shot) - Coins (quarters or dimes) | 2 | Each |
| Squirt Bottle with DI Water | 1 | Each |
| Lab wipes | 1 | Box |

- Safety

Wear lab PPE including lab coat, gloves, and safety glasses.

#### Procedure

  - Part 1: Measuring the Density of Common Liquids.

    - Obtain a 50 ml graduated cylinder

    - Place graduated cylinder onto a top load balance, record the weight, and then tare to scale to zero with the graduated cylinder in place.

Graduated Cylinder empty weight = 	g

    - Obtain sample liquids and record what the liquids are: Liquid #1

Liquid # 2

    - Add approximately 50ml of the liquid to be measured to the graduated cylinder, top off with a pipette to get an accurate volume.

    - Record weight of the liquid (50ml), in the space below

| Document No.: BioWork-3080 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 3 of 4 |
| --- | --- | --- | --- |
| Measuring Density |  |  |  |

    - Calculate the density of the liquid by dividing the weight by its volume. Record in the space below

    - Repeat the steps above for the second liquid and record in the table below.

Note: 1 ml is the same as 1 cc (cubic centimeter); densities are recorded in grams per cubic centimeter (g/cc)

| Name of Liquid Sample | Mass in Grams (of liquid) | Calculate Density in g/cc (d = m/v) | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- |

Note: Calculate the density by dividing the mass in grams by 50 mL (50 cc) and record the density in grams per milliliter (g/cc) in the chart.

  - Part 2 and 3: Measuring the density of common solids

    - Obtain a 250 ml graduated cylinder and a large weigh boat

    - Obtain 2 samples of solid materials and record what they are below: Solid Sample #1

Solid Sample # 2

    - Place large weigh boat onto a top load balance, record the weight, and then tare to scale to zero with the graduated cylinder in place.

Weigh boat empty weight = 	g

    - Record weight of the solid, in the table below

    - Add approximately 100ml of DI water to the 250 ml graduated cylinder, top off with a pipette to get an accurate volume.

    - Add the first solid sample (previously weighed) to the graduated cylinder with the 100 ml of DI water.

    - Record the volume displacement in the table below by subtracting the new volume reading (final) from the volume reading before the solid was added (initial).

    - Calculate density by dividing the total weight of the solid added by the volume displaced (in cc). Record density in table below.

    - Repeat steps for the second solid sample.

| Document No.: BioWork-3080 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 4 of 4 |
| --- | --- | --- | --- |
| Measuring Density |  |  |  |

| Sample # | Description | Weight (g) | Initial Volume (cc) | Final Volume (cc) | Displaced Volume (cc) | Calculated density |
| --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |
| Performed By / Date: |  |  |  |  |  |  |
| Verified By / Date: |  |  |  |  |  |  |

Note: 1 ml is the same as 1 cc (cubic centimeter); densities are recorded in grams per cubic centimeter (g/cc)

#### Notes/Observations

- Attachments

N/A

| Document No.: BioWork-3100 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 1 of 5 |
| --- | --- | --- | --- |
| Working with Proteins |  |  |  |

#### Purpose

To see how factors—such as temperature and pH—affect protein structure, you will study three examples of changes in proteins—two as a result of heat and one as a result of the wrong pH.

#### Scope

This activity consists of three parts.

- Part 1: Denaturing Proteins with Heat

- Part 2: Denaturing Proteins with the Wrong pH

- Part 3: Denaturing an Enzyme with Heat

Note: This is the note section.

#### Responsibilities

Technicians will operate equipment under the direction of an instructor. It is the responsibility of students undergoing training to comply with this procedure.

#### Definitions

Casein:	the protein in milk

#### References

- N/A

#### Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Dissecting Microscope (with under light stage), or light box, or side lighting from a flashlight | 1 | Each |
| Egg White | 1 | Each |
| Beaker | 1 | Each |
| Hot Plate | 1 | Each |
| Petri Dish bottom | 1 | Each |
|  | 1 | Each |

| Document No.: BioWork-3100 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 2 of 5 |
| --- | --- | --- | --- |
| Working with Proteins |  |  |  |

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| O.5 M HCl | 10-20 | Drops |
| Skim Milk |  |  |
| pH Paper | 1 | Roll |
| Measuring Teaspoon (1/2 tsp measure) | 1 | Each |
| Timer | 1 | Each |
| Stir Rod | 1 | Each |
| Hydrogen Peroxide (3% sold in drug store) | 1 | Bottle |
| Slice of Raw Potato | 1 | Each |
| Slice of Cooked Potato | 1 | Each |

- Safety

Wear lab PPE including lab coat, gloves, and safety glasses.

#### Procedure

  - Part 1: Denaturing Proteins with Heat

    - Egg white (albumin) is made mostly of protein. In its natural state, albumin is a clear, gooey liquid.

#### Notes/Observations

| How do egg whites change when they are subjected to high temperatures (boiled or fried)? |  |
| --- | --- |
| Can this be reversed? |  |
| When you cool a cooked egg white, does it return to its former state? |  |

| Document No.: BioWork-3100 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 3 of 5 |
| --- | --- | --- | --- |
| Working with Proteins |  |  |  |

Observe as your instructor uses a pipet to add egg whites—one drop at a time— to a beaker of simmering water. Then answer the following questions.

#### Notes/Observations

| 1. What does the egg white look like after it drops into the hot water? |  |
| --- | --- |
| 2. What is the solid material floating in the beaker of water? |  |

  - Part 2: Denaturing Proteins with the Wrong pH Note:	Milk has a high protein content.

    - Pour skim milk into the petri dish so that the surface of the petri dish is barely covered.

    - Use pH paper to check the pH of the milk. Hold the petri dish over a light source or use a flashlight to side light the dish, and look through the milk. Record the pH and your observation of the milk here.

#### Notes/Observations

    - Use a pipet to add 3 drops of HCl to the milk. Carefully mix the milk solution with a stirring rod. Hold the petri dish over a light source or use a flashlight to side light the dish, and look through the milk. Record your observation of the milk here.

#### Notes/Observations

| Document No.: BioWork-3100 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 4 of 5 |
| --- | --- | --- | --- |
| Working with Proteins |  |  |  |

Repeat steps 8.2.4 until the liquid turns clear. What does the precipitate look like when the milk turns clear?

#### Notes/Observations

    - What is the pH of the milk once it turns clear?

    - How many times did you repeat step 8.2.4 in order to obtain these results?

#### Part 3: Denaturing an Enzyme with Heat

    - In this part of the activity, you will observe how heating destroys the activity of an enzyme in potatoes. The enzyme—called peroxidase—splits hydrogen peroxide into water and oxygen. Peroxides are toxic molecules and peroxidase breaks them down to render them harmless.

#### The reaction that peroxidase catalyzes is: 2 H2O2 → 2 H2O + O2

    - Pour just enough hydrogen peroxide over the raw potato slice to cover it.

    - Watch for a few minutes and record your observations.

Notes and observations:

| Document No.: BioWork-3100 |  | Revision No.: 00 | Effective Date: 26 Aug 2019 |  | Page 5 of 5 |
| --- | --- | --- | --- | --- | --- |
|  | Working with Proteins |  |  |  |  |

    - Microwave a new slice of raw potato for approximately 1.5 minutes. Alternatively dip the raw potato slice into boiling water for 2-4 minutes.

    - Pour just enough hydrogen peroxide over the cooked potato slice to cover it.

    - What happens? Why do you think this reaction is different from the first one?

#### Notes/Observations

- Attachments

N/A

| Document No.: BioWork-3110 | Revision No.: 01 | Effective Date: 26 Aug 2019 | Page 1 of 3 |
| --- | --- | --- | --- |
| Steam Collapse – Instructor Demonstration |  |  |  |

#### Purpose

To learn more about pressure change, your group will observe steam collapse.

#### Scope

N/A

#### Responsibilities

Technicians will operate equipment under the direction of an instructor. It is the responsibility of students undergoing gowning training to comply with this procedure.

#### Definitions

N/A

#### References (Power of Vacuum Demonstrations)

https://www.youtube.com/watch?v=2WJVHtF8GwI https://www.youtube.com/watch?v=VS6IckF1CM0

- Materials

| Material Description |  | Standard Quantity | Unit |
| --- | --- | --- | --- |
| Empty plastic soda bottle with cap |  | 1 | Each |
| Empty aluminum soda Can |  | 1 | Each |
| Thermal Mitt or Tongs |  | 1 | Each |
| Hot Plate (Stir Plate) |  | 1 | Each |
| Funnel or 25ml Pipette |  | 1 | Each |
| Ice Bath (or bucket) |  | 1 | Each |
| 250 ml beaker |  | 1 | Each |
| 1000ml Pyrex beaker |  | 1 | Each |
| Notes: |  |  |  |
|  | BioWork Lab |  |  |

| Document No.: BioWork-3110 | Revision No.: 01 | Effective Date: 26 Aug 2019 | Page 2 of 3 |
| --- | --- | --- | --- |
| Steam Collapse – Instructor Demonstration |  |  |  |

## Safety

Wear lab PPE including lab coat, gloves, and safety glasses.

- Procedure

# Observation of Steam Collapse Using a Soda Bottle

Note: Must start heating water 30 minutes in advance of the demonstration

    - Prepare a boiling bath by adding approximately 500 ml of water to a 1000ml Pyrex beaker. Place stir bar into water, and place beaker on a hot plate stirrer.

    - Prepare Ice Bath by filling an ice bucket about 1/3 with ice, and adding water to completely cover the ice (fill to 2/3 of the capacity of the bucket).

    - Turn the hot plate temperature control up so the water in the beaker comes to a rolling boil.

    - Pour or pipette about 100 mL of the boiling water into the soda bottle and immediately cap it.

    - Immerse the soda bottle into a bucket of ice water.

    - Observe what happens, explain below, and discuss your answer with the class.

Notes/Observations

| Notes: |  |
| --- | --- |
|  | BioWork Lab |

| Document No.: BioWork-3110 | Revision No.: 01 | Effective Date: 26 Aug 2019 | Page 3 of 3 |
| --- | --- | --- | --- |
| Steam Collapse – Instructor Demonstration |  |  |  |

# Observation of Steam Collapse Using a Soda Can

Note: Use same boiling bath and ice baths prepared in section 8.1

    - Turn the hot plate up so the water in the beaker comes to a rolling boil.

    - Rinse and empty the soft drink can.

    - Transfer about 10 – 20 ml of boiling water to the aluminum soda can, to cover the inside bottom of the can with boiling water.

    - Place can back onto the hot plate until steam begins to appear from the opening in the can.

    - When all the water has boiled and the steam has driven the air from the can, quickly (simultaneously) invert the can and submerge it into the ice bath prepared in section 8.1, alternatively into a deep container of cold water.

    - Observe Collapse

Note: The can should collapse and make a loud noise.

Note: It is important to invert and submerge the can immediately after removing it from the heat source. If too much time elapses, air will refill the can and prevent the implosion.

    - Observe what happens, explain below, and discuss your answer with the class.

Notes and observations:

9.0 Attachments: N/A

| Document No.: BioWork-4010 | Revision No.: 00 | Effective Date: 3 Sep 2019 | Page 1 of 5 |
| --- | --- | --- | --- |
| CHEMICAL ENERGY: ENDOTHERMIC AND EXOTHERMIC REACTIONS |  |  |  |

# Purpose

To show the difference between Endothermic and Exothermic chemical reactions. Temperature change is an indirect indicator of the energy change associated with a chemical reaction.

Part 1: Create an Exothermic reaction and measure the temperature change that results. Part 2: Create an Endothermic reaction and measure the associated temperature change.

# Scope

You will do the same technique with each chemical and get a different result each time. Record the data you get and have your readings and measurements verified by a lab partner.

# Responsibilities

It is the responsibility of students undergoing training to comply with this procedure.

# Definitions

Endothermic:	describes reactions in which energy in the form of heat, electricity, or light must be added to the reactants continuously to form the products

Exothermic:	describes reactions in which the reactants interact to form the products and release energy, usually heat

Energy of activation: the minimum amount of energy required to activate atoms or molecules to a condition in which they can undergo chemical transformation or physical transport

# References

  - Endothermic and Exothermic Reactions, Bozeman Science https://www.youtube.com/watch?v=L-G7pLufXAo

  - Tyler DeWitt Phase Changes, Exothermic or Endothermic https://www.tdwscience.com/phase-changes--exothermic-or-endothermic

| Document No.: BioWork-4010 | Revision No.: 00 | Effective Date: 3 Sep 2019 | Page 2 of 5 |
| --- | --- | --- | --- |
| CHEMICAL ENERGY: ENDOTHERMIC AND EXOTHERMIC REACTIONS |  |  |  |

- Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Test tubes – clean and dry; in a test tube rack | 2 | Each |
| Scoopula for transferring chemical | 2 | Each |
| Thermometer (Digital if available) | 2 | Each |
| Ammonium Nitrate, NH₄NO₃ (Must be kept closed) | 1 | Container |
| Calcium Chloride, CaCl₂ | 1 | Container |
| DI water in squirt bottle | 1 | Each |
| Paper towel | 1 | Roll |

- Safety

Wear fully fastened Lab coat, eye protection and gloves. The result of your reactions may be rinsed down the sink with plenty of water. Clean equipment before and after use by rinsing with DI and drying with a paper towel or lab wipe.

| Document No.: BioWork-4010 | Revision No.: 00 | Effective Date: 3 Sep 2019 | Page 3 of 5 |
| --- | --- | --- | --- |
| CHEMICAL ENERGY: ENDOTHERMIC AND EXOTHERMIC REACTIONS |  |  |  |

# Procedure

Part 1: Exothermic Reaction

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Scoop CaCl₂ to about 1 cm or 3/8 in. depth into the first test tube. Note: If a digital thermometer is available, place probe into empty tube first, then add the chemical. |  |  |
| 2. | Take the temperature, record it below: Temperature °C (Initial) |  |  |
| 3. | Add enough DI water from the squirt bottle to just cover the chemical. Note: CaCl₂ is fine and must be stirred with a pipette or stir rod after the addition of water. |  |  |
| 4. | Take the temperature, record it below: Temperature °C (Post Reaction) |  |  |
| 5. | Feel the test tube with your hand, note any changes you observe: What Happened? |  |  |

| Document No.: BioWork-4010 | Revision No.: 00 | Effective Date: 3 Sep 2019 | Page 4 of 5 |
| --- | --- | --- | --- |
| CHEMICAL ENERGY: ENDOTHERMIC AND EXOTHERMIC REACTIONS |  |  |  |

Part 2: Endothermic Reaction

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Scoop NH₄NO₃ to about 1 cm or 3/8 in. depth into the second test tube. Note: If a digital thermometer is available, place probe into empty tube first, then add the chemical. |  |  |
| 2. | Take the temperature, record it below: Temperature °C (Initial) |  |  |
| 3. | Add enough DI water from the squirt bottle to just cover the chemical. If necessary, carefully stir the solution. |  |  |
| 4. | Take the temperature, record it below: Temperature °C (Post Reaction) |  |  |
| 5. | Feel the test tube with your hand, note any changes you observe: What Happened? |  |  |

# Figures and Diagrams

N/A

- Notes/Observations

Which of these chemicals could be used to melt ice on the road?

## Which could be used in place of the salt in a home ice cream maker?

Did either reaction require Energy of Activation?

# Attachments

N/A

| Document No.: BioWork-4030 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 1 of 4 |
| --- | --- | --- | --- |
| Getting information from a SDS |  |  |  |

#### Purpose

To practice finding safety information in a Safety Data Sheet.

#### Scope

This activity will look at several SDS sheets for various compounds.

#### Responsibilities

Technicians will work in the lab under the direction of an instructor. It is the responsibility of students undergoing training to comply with this procedure.

#### Definitions

SDS:	Safety Data Sheet

#### References

- Safety Data Sheet: Sodium Sulfite, Anhydrous http://www.labchem.com/tools/msds/msds/LC24930.pdf

- Safety Data Sheet: Methanol https://beta-

static.fishersci.com/content/dam/fishersci/en_US/documents/programs/education/regulatory- documents/sds/chemicals/chemicals-m/S25426A.pdf

- Safety Data Sheet: Sulfuric Acid https://beta-

static.fishersci.com/content/dam/fishersci/en_US/documents/programs/education/regulatory- documents/sds/chemicals/chemicals-s/S25899.pdf

- Safety Data Sheet: Potassium Hydroxide https://www.fishersci.com/msdsproxy%3FproductName%3DP2503%26productDescription% 3DPOT%2BHYDROXIDE%2BCERT%2BACS%2B3KG%26catNo%3DP2503%26vendorId

%3DVN00033897%26storeId%3D10652

| Document No.: BioWork-4030 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 2 of 4 |
| --- | --- | --- | --- |
| Getting information from a SDS |  |  |  |

#### Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| SDS For Sodium Sulfite | 1 | Each |
| SDS For Methanol | 1 | Each |
| SDS For Sulfuric Acid | 1 | Each |
| SDS For Potassium Hydroxide | 1 | Each |

- Safety

Wear lab PPE including lab coat, gloves, and safety glasses.

#### Procedure

  - Use information from the SDS to answer the questions that follow.

    - Sodium Sulfite

| What does sodium sulfite look like? |  |
| --- | --- |
| What type of a health hazard is it? |  |
| Is it possible that you might be allergic to sodium sulfite? |  |

    - Methanol

| Is it dangerous to ingest methanol? |  |
| --- | --- |
| What should you do if one of your classmates drinks methanol? |  |

| Notes: |  |
| --- | --- |
|  | BioWork Lab |

| Document No.: BioWork-4030 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 3 of 4 |
| --- | --- | --- | --- |
| Getting information from a SDS |  |  |  |

    - Sulfuric Acid

| What are the hazards for getting sulfuric acid on your skin? |  |
| --- | --- |
| What should you do if you get sulfuric acid on your skin? |  |
| What PPE should you wear when handling concentrated sulfuric acid? |  |
| How should you store sulfuric acid? |  |

    - Potassium Hydroxide

| How should you dispose of unused potassium hydroxide? |  |
| --- | --- |
| Is the chemical a fire and/or explosion hazard? If yes, which one? |  |
| What is the health rating for the chemical? |  |
| Provide an overview of the health hazards. |  |
| Below the airborne exposure limit, what PPE should you wear when working with the KOH? |  |

| Document No.: BioWork-4030 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 4 of 4 |
| --- | --- | --- | --- |
| Getting information from a SDS |  |  |  |

#### Figures and Diagrams

N/A

#### Notes/Observations

See above

#### Attachments

N/A

| Document No.: BioWork-4060 | Revision No.: 02 | Effective Date: 22 Aug 2019 | Page 1 of 7 |
| --- | --- | --- | --- |
| Making Media |  |  |  |

#### Purpose

To learn what making media involves, you will:

- Make a simple type of media called LB+. LB+ is a variation on a type of media commonly used for growing E. coli bacteria.

- Add various ingredients at different points in the process to see if there is any effect on how the media turns out.

#### Scope

This activity has two parts:

Part 1: Making the Media for Autoclaving

Part 2: Observing Effects of Different Protocols

Note: This is the note section.

#### Responsibilities

It is the responsibility of students undergoing training to comply with this procedure. With your team, look at the tasks in the activity and decide who will do what. Make sure one team member is assigned to double-check all measurements.

#### Definitions

Lysogeny broth (LB):	is a nutritionally rich medium primarily used for the growth of bacteria. Its creator, Giuseppe Bertani, intended LB to stand for lysogeny broth, but LB has also come to be commonly referred to as Luria broth, Lennox broth, or Luria-Bertani medium.

#### References: N/A

| Document No.: BioWork-4060 | Revision No.: 02 | Effective Date: 22 Aug 2019 | Page 2 of 7 |
| --- | --- | --- | --- |
| Making Media |  |  |  |

- Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Balance | 1 | Each |
| pH meter | 1 | Each |
| Magnetic stirrer/hot plate | 1 | Each |
| Stirring bar | 1 | Each |
| Weigh boats | 3 | Each |
| 250-mL beaker | 1 | Each |
| 100-mL graduated cylinder | 1 | Each |
| 50-mL bottles with caps | 4 | Each |
| 5-mL pipettes | 4 | Each |
| 1-mL pipettes | 9 | Each |
| Pipette aid | 1 | Each |
| Tryptone | 1 | gram |
| Yeast extract | 0.5 | gram |
| NaCl | 0.05 | gram |
| NaOH, 1M | 1 | mL |
| 65% glucose | 20 | mL |
| CaCl2 (calcium chloride), 2M | 5 | mL |
| MgSO4▪7H20 (magnesium sulfate), 2M | 5 | mL |
| Squirt bottle of DI water | 1 | Each |

| Document No.: BioWork-4060 | Revision No.: 02 | Effective Date: 22 Aug 2019 | Page 3 of 7 |
| --- | --- | --- | --- |
| Making Media |  |  |  |

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Aluminum foil | 1 | Each |
| Marking pen | 1 | Each |
| Label tape | 1 | Each |

#### Safety

Wear lab PPE including lab coat, gloves, and safety glasses.

#### Procedure

Activity and/or process

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Weigh out the following solid ingredients and put them in the beaker: 1.0 g tryptone 0.5 g yeast extract 0.05 g NaCl Record actual weights below: 1.0 g tryptone: g 0.5 g yeast extract: g 0.05 g NaCl: g |  |  |
| 2. | If any ingredients get stuck on the side, wash them down into the bottom with water from a squirt bottle. |  |  |
| 3. | Add 80 mL DI water; add a stirring bar, put the beaker on a magnetic stirrer, and make sure all the ingredients get dissolved in the water |  |  |
| 4. | Check pH and adjust it to 7.0 with 0.2 mL 1M NaOH if necessary. |  |  |

| Document No.: BioWork-4060 | Revision No.: 02 | Effective Date: 22 Aug 2019 | Page 4 of 7 |
| --- | --- | --- | --- |
| Making Media |  |  |  |

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 5. | Transfer medium to a graduated cylinder and bring final volume up to 100 mL with DI water. |  |  |
| 6. | Put 25 mL of media in each of the four bottles. Cover the outside of each bottle with foil but do not cover the cap. Use a piece of label tape on the foil to label the bottles 1, 2, 3, and 4 and place a piece of autoclave tape on the caps. |  |  |
| 7. | Follow the procedures below for each bottle: Bottle 1: Loosely cap Bottle 2: Add 4 mL of 65% glucose; Loosely cap Bottle 3: Add 0.8 mL of 2M CaCl2 (calcium chloride) Add 0.4 ml of 2M MgSO4▪7H20 (magnesium sulfate); Loosely cap Bottle 4 Add 4 mL of 65% glucose; Add 0.8 mL of 2M CaCl2 (calcium chloride); and 0.4 ml of 2M MgSO4▪7H20 (magnesium sulfate); Loosely cap |  |  |
| 8. | Autoclave and cool all four bottles. Note: Autoclave must be run for liquid cycle (slow exhaust); allow approximately 2 hours for cycle completion and cool down |  |  |
| 9. | Note: Your instructor will provide you with separately sterilized solutions of glucose, calcium chloride, and magnesium sulfate. After the autoclaved bottles have cooled enough to handle, uncap them but do not remove the outer foil. Then carry out the procedures described below for each bottle. |  |  |

| Document No.: BioWork-4060 | Revision No.: 02 | Effective Date: 22 Aug 2019 | Page 5 of 7 |
| --- | --- | --- | --- |
| Making Media |  |  |  |

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 10 | Bottle 1: Add 4 mL of 65% glucose; Add 0.8 mL of 2M CaCl2 (calcium chloride); and 0.4 ml of 2M MgSO4▪7H20 (magnesium sulfate); Swirl the bottle to mix the ingredients Bottle 2: Add 0.8 mL of 2M CaCl2 (calcium chloride) Add 0.4 ml of 2M MgSO4▪7H20 (magnesium sulfate); Swirl the bottle to mix the ingredients Bottle 3: Add 4 mL of 65% glucose; Swirl the bottle to mix the ingredients Bottle 4 Swirl the bottle to mix the ingredients |  |  |
| 11 | Note: All four bottles now contain exactly the same ingredients. The only difference is the order in which the ingredients were added. |  |  |
| 12 | Unwrap and inspect the bottles. Answer questions listed below in the notes/Observations section. |  |  |

| Document No.: BioWork-4060 | Revision No.: 02 | Effective Date: 22 Aug 2019 | Page 6 of 7 |
| --- | --- | --- | --- |
| Making Media |  |  |  |

#### Figures and Diagrams

Flow Diagram for LB+ Procedure:

| Document No.: BioWork-4060 | Revision No.: 02 | Effective Date: 22 Aug 2019 | Page 7 of 7 |
| --- | --- | --- | --- |
| Making Media |  |  |  |

#### Notes/Observations

  - Does the medium look the same in all four?

  - Is there any browning or precipitation in any of the bottles?

  - Record the appearance of the media in each bottle on the observation form that follows.

  - The only difference in the bottles will have been the order in which some of the ingredients have been added during the process.

| Observations |  |  |  |
| --- | --- | --- | --- |
| Bottle 1 | Bottle 2 | Bottle 3 | Bottle 4 |

#### Making Growth Media: Key Concepts 8

    - The first concept can be worded "Making a mistake in the media recipe could adversely affect the product.

    - The order of adding ingredients may be critical.

    - All the media must actually get down into the tank and get mixed up.

    - Use your eyes and nose to check the finished mixture (ingredients are correct, everything is dissolved, there is no precipitation or browning, smells are normal).

    - Many media components, especially for mammalian cell culture, are heat- sensitive and must be filter-sterilized.

    - Media preparation areas must be kept clean and spills should be cleaned up immediately.

    - Follow the SOP exactly and always be consistent.

    - Have a coworker double-check your measurements.

| Document No.: BioWork-4075 | Revision No.: 00 | Effective Date: 11 Nov 2019 | Page 1 of 8 |
| --- | --- | --- | --- |
| Making and Observing the Characteristics of a Buffer |  |  |  |

#### Purpose

A buffer solution is a solution of, usually, a weak acid and its conjugate base, or, less commonly, a weak base and its conjugate acid. A buffer solution resists changes in the concentrations of the hydronium ion and hydroxide ion (and therefore pH) when the solution is diluted or when small amounts of an acid or base are added to it.

One common example of a buffer solution is a solution of acetic acid (the weak acid) and sodium acetate (its conjugate base). In solution, acetic acid reaches an equilibrium illustrated by the following equation.

CH3COOH(aq) + H2O(I) ⇌ CH3COO– + H3O+

#### Scope

In this lab, we’ll make up a buffer solution of acetic acid and sodium acetate and examine the effects of adding hydrochloric acid and sodium hydroxide to this buffer solution.

#### Responsibilities

It is the responsibility of students undergoing training to comply with this procedure.

#### Definitions

Buffer:	A buffer is a solution that can resist pH change upon the addition of an acidic or basic components. It is able to neutralize small amounts of added acid or base, thus maintaining the pH of the solution relatively stable

#### References

  - MIT Open Courseware: Buffers

https://ocw.mit.edu/resources/res-tll-004-stem-concept-videos-fall-2013/videos/structure- function- properties/buffers/

| Document No.: BioWork-4075 | Revision No.: 00 | Effective Date: 11 Nov 2019 | Page 2 of 8 |
| --- | --- | --- | --- |
| Making and Observing the Characteristics of a Buffer |  |  |  |

#### Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| PPE: goggles, gloves, and protective clothing | 1 | Per Person |
| Beaker, 150 mL | 2 | Each |
| Beaker, 250 mL | 2 | Each |
| Graduated cylinder, 100 mL | 1 | Each |
| Graduated cylinder, 10 mL | 1 | Each |
| Pipette, 1.00 mL with tips, or 1mL transfer pipette | 1 | Each |
| Stirring Rod or Magnetic Stir Plate and stir bars | 1 | Each |
| pH Meter (or pH paper if no meter available) | 1 | Each |
| Acetic acid, 1.0 M (100 mL) | 100 | mL |
| Sodium acetate, 1.0 M (100 mL) | 100 | mL |
| Hydrochloric acid, 6.0 M (20 mL) | 20 | mL |
| Sodium hydroxide, 6.0 M (20 mL) | 20 | mL |
| Distilled or deionized water | 100 | mL |
| Top Load Balance (0-210g) or equivalent | 1 | Each |
| Universal Indicator (optional) | 1 | As needed |

- Safety

Wear lab PPE including lab coat, gloves, and safety glasses when working in the lab.

CAUTION:	Hydrochloric acid and sodium hydroxide are corrosive. Wear splash goggles, gloves, and protective clothing at all times.

| Document No.: BioWork-4075 | Revision No.: 00 | Effective Date: 11 Nov 2019 | Page 3 of 8 |
| --- | --- | --- | --- |
| Making and Observing the Characteristics of a Buffer |  |  |  |

#### Procedure

General Notes: Preparation

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Make up 1.0 M acetic acid by diluting 5.7 mL of concentrated (glacial) acetic acid to 100 mL. Label the beaker. Alternatively, you may use distilled white vinegar, which is close to 1 M straight out of the bottle. |  |  |
| 2. | Make 100 mL of 1.0 M sodium acetate by dissolving 8.20 g of anhydrous sodium acetate or 13.61 g of sodium acetate trihydrate in some distilled water and then making up the solution to 100 mL. Label the beaker. |  |  |
| 3. | Make 20 mL of 6.0 M hydrochloric acid by mixing 10.00 mL of concentrated (37%, 12 M) hydrochloric acid with 10.00 mL of water. Label the beaker. Alternatively use hardware-store muriatic acid (31.45%, 10.3 M), mix 11.65 mL of the acid with 8.35 mL of water Safety Note: Add acid to water! (NOT water to acid) |  |  |
| 4. | Make 20 mL of 6.0 M sodium hydroxide by dissolving 4.80 g of sodium hydroxide, stirring constantly, in 90 mL of water and making up the solution to 100 mL. Label the beaker. Safety Note: Add NaOH to water! (NOT water to NaOH) Caution: this reaction is extremely exothermic. |  |  |

| Document No.: BioWork-4075 | Revision No.: 00 | Effective Date: 11 Nov 2019 | Page 4 of 8 |
| --- | --- | --- | --- |
| Making and Observing the Characteristics of a Buffer |  |  |  |

Procedure

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | If you have not already done so, put on your splash goggles, gloves, and protective clothing |  |  |
| 2. | Make up the buffer solution by mixing 100 mL of 1.0 M acetic acid and 100 mL of 1.0 M sodium acetate in the 250 mL beaker. Label the beaker. |  |  |
| 3. | Transfer 100 mL of the buffer solution to one of the 150 mL beakers, and 100 mL of distilled water to the second 150 mL beaker. Label all beakers. |  |  |
| 4. | Read and follow the directions for your pH meter with respect to calibrating it, rinsing the electrode between measurements and other handling operations. |  |  |
| 5. | Use the pH meter or pH test paper to measure the pH of the buffer solution, and record the observed value on Line A in the second and fourth columns of the observation Table in section 10. |  |  |
| 6. | Use the pH meter to measure the pH of the distilled water, and record the observed value on Line A in the third and fifth columns of the observation Table in section 10. |  |  |
| 7. | Use the 1.00 mL pipette to transfer 1.00 mL of 6.0 M hydrochloric acid to the beaker that contains the buffer solution and another 1.00 mL of hydrochloric acid to the beaker that contains the distilled water. Stir or swirl the beakers to mix the solutions thoroughly. Note: If you use only one stirring rod, rinse it thoroughly before using it in the other beaker. |  |  |
| 8. | Use the pH meter to determine the pH value for the buffer solution, and record that value on Line B in the second column of the observation Table in section 10. |  |  |
| 9. | Use the pH meter to determine the pH value for the water solution, and record that value on Line B in the third column of the observation Table in section 10. |  |  |
| 10. | Repeat steps 7 through 9, adding 1.00 mL of hydrochloric acid each time until you have added a total of 10.00 mL of hydrochloric acid to the buffer solution and water solution. |  |  |

| Document No.: BioWork-4075 | Revision No.: 00 | Effective Date: 11 Nov 2019 | Page 5 of 8 |
| --- | --- | --- | --- |
| Making and Observing the Characteristics of a Buffer |  |  |  |

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 11. | Rinse the beakers and pipette thoroughly. |  |  |
| 12. | Transfer the remaining 100 mL of buffer solution to one of the 150 mL beakers, and 100 mL of distilled water to the second 150 mL beaker. |  |  |
| 13. | Repeat steps 7 through 11 using the 6.0 M sodium hydroxide solution. |  |  |
| 14. | Compare the results from the Observation Table and comment on the effect of buffer vs. water in resisting pH change. |  |  |
| 15. | Disposal: All of the solutions used in this laboratory can be flushed down the drain with plenty of water. |  |  |

Other Notes (Optional)

  - Phenolphthalein is an indicator which usually is colorless. Add 3 drops of phenolphthalein to 100 ml of water, you will obtain a colorless solution. Solution should stay colorless in an acid, but turn red in a base.

  - Also can use Universal Indicator which is colored.

| Document No.: BioWork-4075 | Revision No.: 00 | Effective Date: 11 Nov 2019 | Page 6 of 8 |
| --- | --- | --- | --- |
| Making and Observing the Characteristics of a Buffer |  |  |  |

#### Figures and Diagrams

Prep Diagram

1.0 M acetic acid

1.0 M sodium

acetate

100 mL

100 mL

6.0 M hydrochloric

acid

20 mL

Add in 1ml

increments, checking pH at each addition

Buffer Solution

100 mL

Buffer Solution

6.0 M sodium

hydroxide

20 mL

DI Water

Add in 1ml increments, checking pH at each addition

100 mL

Buffer Solution

DI Water

| Document No.: BioWork-4075 | Revision No.: 00 | Effective Date: 11 Nov 2019 | Page 7 of 8 |
| --- | --- | --- | --- |
| Making and Observing the Characteristics of a Buffer |  |  |  |

Notes/Observations

Table: Observe the characteristics of a buffer solution – observed data

| Acid/Base Added | Buffer + HCI | Water + HCI | Buffer + NaOH | Water + NaOH |
| --- | --- | --- | --- | --- |
| A. 0.00 mL | . pH | . pH | . pH | . pH |
| B. 1.00 mL | . pH | . pH | . pH | . pH |
| C. 2.00 mL | . pH | . pH | . pH | . pH |
| D. 3.00 mL | . pH | . pH | . pH | . pH |
| E. 4.00 mL | . pH | . pH | . pH | . pH |
| F. 5.00 mL | . pH | . pH | . pH | . pH |
| G. 6.00 mL | . pH | . pH | . pH | . pH |
| H. 7.00 mL | . pH | . pH | . pH | . pH |
| I. 8.00 mL | . pH | . pH | . pH | . pH |
| J. 9.00 mL | . pH | . pH | . pH | . pH |
| K. 10.00 mL | . pH | . pH | . pH | . pH |

| Document No.: BioWork-4075 | Revision No.: 00 | Effective Date: 11 Nov 2019 | Page 8 of 8 |
| --- | --- | --- | --- |
| Making and Observing the Characteristics of a Buffer |  |  |  |

#### Attachments

Example pH Response

|  | Acid Additions (pH) |  | Base Additions (pH) |  |
| --- | --- | --- | --- | --- |
| Acid/Base Added (mL) | Acetate Buffer | DI Water | Acetate Buffer | DI Water |
| 0 | 4.66 | 6.31 | 4.66 | 6.31 |
| 1 | 4.62 | 1.48 | 4.75 | 12.55 |
| 2 | 4.41 | 1.12 | 4.88 | 12.81 |
| 3 | 4.42 | 0.95 | 5.02 | 12.95 |
| 4 | 4.30 | 0.83 | 5.17 | 13.05 |
| 5 | 4.18 | 0.74 | 5.37 | 13.12 |
| 6 | 4.06 | 0.65 | 5.66 | 13.18 |
| 7 | 3.90 | 0.58 | 6.30 | 13.24 |
| 8 | 3.74 | 0.53 | 12.26 | 13.28 |
| 9 | 3.53 | 0.48 | 12.67 | 13.33 |
| 10 | 3.24 | 0.43 | 12.86 | 13.38 |

| Document No.: BioWork-4095 | Revision No.: 00 | Effective Date: 22 Aug 2019 | Page 1 of 10 |
| --- | --- | --- | --- |
| Working with pH |  |  |  |

#### Purpose

Measuring Process Variables i.e. measuring and controlling temperature, pressure, flow, and level is an important part of a process technician’s work. Another important process variable is pH. In this activity, you will practice skills needed to measure and adjust pH.

#### Scope

This activity has 4 parts:

Part 1: Changing Color with pH Part 2: Calibrating the pH Meter Part 3: Measuring pH

Part 4: Changing the pH of Water

#### Responsibilities

It is the responsibility of students undergoing training to comply with this procedure.

#### Definitions

pH:	pH is a measure of hydrogen ion concentration, a measure of the acidity or alkalinity of a solution.

The pH scale usually ranges from 0 to 14. Aqueous solutions at 25°C with a pH less than 7 are acidic, while those with a pH greater than 7 are basic or alkaline.

#### References

  - Tyler DeWitt, Chemistry, pH video (Tricks with Acids and Base) https://www.tdwscience.com/a-colorful-magic-trick-with-acids-and-ba

  - Bio-Rad Video, How to Use a pH Meter https://www.youtube.com/watch?v=vwY-xWMam7o

#### Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| PART 1 |  |  |
| NuGrape® soda | 1 | 12-16oz |
| 50mL beakers | 2 | each |
| Full-range pH paper | 1 | roll |

| Document No.: BioWork-4095 | Revision No.: 00 | Effective Date: 22 Aug 2019 | Page 2 of 10 |
| --- | --- | --- | --- |
| Working with pH |  |  |  |

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Gloves | 2 | pairs |
| 10mL pipettes | 2 | each |
| Pipet-Aid® | 1 | each |
| 4N NaOH | NS | NS |
| PART 2 |  |  |
| pH 7.0 standard solution | NS | NS |
| pH 4.0 standard solution | NS | NS |
| pH meter | 1 | each |
| Magnetic stirrer/hot plate and Magnetic stir bar | 1 | each |
| Ring stand and Clamp for electrode | 1 | each |
| 500 mL Beaker | 1 | each |
| Lab Wipes | 1 | each |
| Label Tape and Marking Pen | 1 | each |
| Squirt Bottle of DI Water | 1 | each |
| PART 3 |  |  |

| Document No.: BioWork-4095 | Revision No.: 00 | Effective Date: 22 Aug 2019 | Page 3 of 10 |
| --- | --- | --- | --- |
| Working with pH |  |  |  |

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Materials to be tested: Examples: - Vinegar - Soft drink - Orange juice or Pineapple Juice - Milk - DI water - Milk of Magnesia - Household ammonia - Window Cleaner - Bleach - Other materials collected earlier | 10-30 mL | Each |
| PART 4 |  |  |
| 100 mL Beakers | 2 | Each |
| 50mL Beakers | 2 | Each |
| 50 mL Graduated Cylinder | 1 | Each |
| 4N HCl (Hydrochloric Acid) | NS | NS |
| Magnetic stirrer/hot plate and Magnetic stir bar | 1 | Each |
| Magnetic stir bar retriever | 1 | Each |
| Disposable transfer pipettes | 4 | Each |
| Other materials collected earlier (Part 2 & 3) | NS | NS |

| Document No.: BioWork-4095 | Revision No.: 00 | Effective Date: 22 Aug 2019 | Page 4 of 10 |
| --- | --- | --- | --- |
| Working with pH |  |  |  |

- Safety

During this activity, you will use a strong acid (HCl) and a strong base (NaOH). Observe the appropriate safety precautions as directed by your instructor.

- You will also use some common household acids and bases. While they are not as hazardous, you should still observe the same precautions when you work with them during the activity.

- You must wear safety glasses and lab coats through the entire activity. Make sure all snaps or buttons are fastened.

- When handling concentrated acid and base solutions, wear latex gloves—or nitrile gloves if you are allergic to latex.

- When working with HCl, keep the bottle capped except when pipetting small volumes from it. As you have learned, HCl produces dangerous fumes.

- If any spill of concentrated acid or base occurs, notify your instructor immediately. All safety information here.

| Document No.: BioWork-4095 | Revision No.: 00 | Effective Date: 22 Aug 2019 | Page 5 of 10 |
| --- | --- | --- | --- |
| Working with pH |  |  |  |

Procedure

Part 1: Changing Color with pH

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | During this activity, you will work with your partner to see how the pH of a solution can affect the nature of a chemical in the solution. In this case, changing the pH of a solution causes the solution to change its color. | N/A | N/A |
| 2. | Collect first 7 items from the material list: Soda, beakers, pH paper, gloves, pipettes, pipet-Aid, 4N NaOH |  |  |
| 3. | Observe as your instructor demonstrates the use of Pipet- Aids® when pipetting strong acids and bases. |  |  |
| 4. | Use a pipette to measure out 10mL of NuGrape® soda into a 50mL beaker. |  |  |
| 5. | Tear off about 11⁄2 inches of full-range pH paper and use it to measure the pH of the soda. Record initial pH |  |  |
| 6. | Pour about 5mL of 4N NaOH into a beaker and label the beaker. This is done to prevent contamination that can occur if you pipette directly from the bottle. |  |  |
| 7. | Pipette 3mL of NaOH into the beaker of NuGrape® soda and swirl to mix. |  |  |
| 8. | Observe any changes in the appearance of the soda and record. Record Changes: |  |  |
| 9. | Use another 11⁄2-inch strip of pH paper to check the pH of the soda now. Record the pH after adjustment |  |  |

| Document No.: BioWork-4095 | Revision No.: 00 | Effective Date: 22 Aug 2019 | Page 6 of 10 |
| --- | --- | --- | --- |
| Working with pH |  |  |  |

Part 2: Calibrating the pH Meter

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | The procedure in this activity is called a two-point calibration because two pH standards (7.0 and 4.0) are used. You will check to see that the meter accurately measures a solution with a known pH of 7.0 and a solution with a known pH of 4.0. | N/A | N/A |
| 2. | Watch as your instructor demonstrates how to set up the pH meter. Then use the instructions that follow to set up your pH meter. |  |  |
| 3. | Clamp the electrode to the ring stand, as demonstrated by your instructor. Check that the electrode leads are plugged into the meter. |  |  |
| 4. | Open the filling hole above the body of the electrode. When the electrode is in use, the filling hole should always be in the open position. |  |  |
| 5. | Remove the plastic cap from the electrode bulb. |  |  |
| 6. | Rinse the electrode. With the 500mL beaker positioned to catch the rinse water, use the squirt bottle to rinse the electrode with DI water. Note: For the rest of this activity, you will be using the 500mL beaker to catch the rinse water. |  |  |
| 7. | Blot—do not rub—the end of the electrode with a laboratory wipe. |  |  |
| 8. | Press the on/off button to turn on the pH meter. Press the mode button to access the Measure screen. |  |  |
| 9. | Press the setup button and then the enter button to clear the meter of any previous readings. Standardize (calibrate) at pH 7.0; Press the std button to view the Standardization screen. |  |  |
| 10. | Immerse the electrode in the pH 7.0 standard solution. Press the std button again to begin standardization. When the pH reading is stable, the meter will return automatically to the Measure screen and display a reading between 6.8 and 7.2. |  |  |

| Document No.: BioWork-4095 | Revision No.: 00 | Effective Date: 22 Aug 2019 | Page 7 of 10 |
| --- | --- | --- | --- |
| Working with pH |  |  |  |

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 11. | Remove the electrode from the solution. Use the squirt bottle to thoroughly rinse the electrode with DI water and blot dry with a wipe. |  |  |
| 12. | Standardize (calibrate) at pH 4.0; Repeat steps 10 through 11—this time with pH 4.0 solution. |  |  |
| 13. | Store Electrode in appropriate storage buffer (pH 7 or manufacturers recommendation). Note: pH meter is calibrated and ready to measure samples |  |  |

Part 3: Measuring pH of Various Samples

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Before making your pH measurements, estimate the pH range you expect for each sample as neutral (7), acidic (0- 7), or basic (7-14) and enter in the table below. |  |  |
| 2. | Remove the electrode from the storage solution. Use the squirt bottle to rinse the electrode with DI water and blot dry. |  |  |
| 3. | Put the electrode into the solution to be tested and press the mode button on the pH meter, make sure the meter is in pH mode. |  |  |
| 4. | Read the pH measure and record the actual value in the table. |  |  |
| 5. | Remove the electrode from the sample solution. Use the squirt bottle to rinse the electrode with DI water and blot with a lab wipe. |  |  |
| 6. | Repeat steps 3 through 5 for all samples. |  |  |
| 7. | When finished, use the squirt bottle to rinse the electrode with DI water, blot, and return it to the storage solution. |  |  |
| 8. | Compare estimated value with actual pH |  |  |

| Document No.: BioWork-4095 | Revision No.: 00 | Effective Date: 22 Aug 2019 | Page 8 of 10 |
| --- | --- | --- | --- |
| Working with pH |  |  |  |

Part 4: Changing the pH of Water (to 10)

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Measure 50mL of DI water from your squirt bottle into a 100mL beaker. Use the procedure you learned in Part 3 to check the pH of the water and record it. pH before adjustment |  |  |
| 2. | Answer these questions: - To adjust the DI water to a pH of 10, will you add acid or base? - . How much do you think you will need to add? |  |  |
| 3. | Use the squirt bottle to rinse the magnetic stir bar with DI water and place it in the beaker. |  |  |
| 4. | Place the beaker on the magnetic stirrer/hot plate. |  |  |
| 5. | Remove the electrode from the storage solution. |  |  |
| 6. | Use the squirt bottle to thoroughly rinse the electrode with DI water and blot dry with a wipe. |  |  |
| 7. | Lower the electrode into the beaker, making sure that the electrode bulb is under the surface of the liquid but not touching the stir bar. |  |  |
| 8. | Turn on the magnetic stir bar to the lowest possible setting to agitate the solution. |  |  |
| 9. | Press the mode button and check that the pH meter is in the pH mode. |  |  |
| 10. | Decide whether to add an acid (4N HCl) □ or a base (4N NaOH) □ |  |  |
| 11. | Pour a few milliliters of the acid or base into a disposable cup and label it. |  |  |

| Document No.: BioWork-4095 | Revision No.: 00 | Effective Date: 22 Aug 2019 | Page 9 of 10 |
| --- | --- | --- | --- |
| Working with pH |  |  |  |

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 12. | Use a disposable transfer pipette to add the acid or base – one drop at a time – to the DI water. Important: Be careful. It is easy to get impatient, add too much, and overshoot the desired pH. |  |  |
| 13. | Allow time for the stir bar to mix the solution and for the pH meter to give you a stable reading. |  |  |
| 14. | Observe the pH meter reading and add more acid or base until the desired pH is reached. |  |  |
| 15. | Answer the questions that follow: - What did you add—the acid or the base? - How many drops did you add? - What was the final pH of the solution? pH after adjustment |  |  |
| 16. | Set the function knob to standby and then remove the electrode from the solution. |  |  |
| 17. | Use the squirt bottle to rinse the electrode with DI water, blot, and return it to the storage solution. |  |  |
| 18. | Clean and dry the magnetic stir bar. |  |  |

#### Figures and Diagrams

Diagram 1: N/A

| Document No.: BioWork-4095 | Revision No.: 00 | Effective Date: 22 Aug 2019 | Page 10 of 10 |
| --- | --- | --- | --- |
| Working with pH |  |  |  |

#### Notes/Observations

Instructor Note: Can make a pH indicator (liquid) by boiling red cabbage leaves for 5 minutes, let sit for 15-20 min, and decant off the blue liquid (indicator); will go pink to red for acids, light blue to green to yellow for bases.

Measuring pH Sample Table (Part 3 results)

| Sample Tested (name) | Expected pH Acidic, Neutral, Basic | Actual pH (measured) |
| --- | --- | --- |
| Done By: Date: Checked By: Date: |  |  |

#### Attachments

N/A

| Document No.: BioWork-5010 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 1 of 6 |
| --- | --- | --- | --- |
| Reading a Process Diagram |  |  |  |

#### Purpose

To practice reading a process diagram, you will study a diagram, work through the process flow and equipment, and answer questions about the process diagram. To complete this activity, you will use a typical process diagram for the water purification system in a biopharmaceutical manufacturing plant as well as lead sheets, also called legends, that show what the abbreviations and symbols in the diagram mean.

#### Scope

Look at the process diagram that follows. Listen to your instructor talk through the key points, specifically:

- The process flows from left to right (black arrowheads indicate the direction of flow in each line).

- To simplify the diagram, not all equipment has been labeled with an equipment number.

- P&IDs do not show scale or the layout of the equipment—only the order of things.

- The input to the process in the diagram is water that has been through water softening and activated charcoal units.

#### Responsibilities

Technicians will operate equipment under the direction of an instructor. It is the responsibility of students undergoing training to comply with this procedure.

#### Definitions

P&ID: A piping and instrumentation diagram (P&ID) is a detailed diagram in the process industry which shows the piping and process equipment together with the instrumentation and control devices.

#### References

- P&ID Symbols https://www.lucidchart.com/pages/p-id-symbols-legend

N/A

#### Materials

- Safety

Wear lab PPE including lab coat, gloves, and safety glasses.

| Document No.: BioWork-5010 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 2 of 6 |
| --- | --- | --- | --- |
| Reading a Process Diagram |  |  |  |

#### Procedure

Note: The unit operations shown in the diagram are:

- Reverse osmosis (RO)

- Ultraviolet light radiation (UV)

- Deionization (CDI)

- Microfiltration (FL)

- The output from this process is WFI (Water for Injection). WFI is sent to a purified water storage tank and is ready for use in the plant.

  - Identify each unit operation on the diagram. Except for filtration, the operations are drawn as simple blocks.

  - Identify duplicate equipment. The process contains duplicate equipment so it can operate continuously. One set of equipment can be taken offline for servicing or cleaning while the other operates, or both sets of equipment can operate together as needed.

  - Trace the entire process flow by following the water through all equipment and into the purified water storage tank. Ignore the duplicate equipment. Trace the flow through only one RO unit, one CDI unit, and one FL unit.

  - Find the single waste stream shown in the diagram. What kind of impurities might be in this waste stream?

  - Find the sampling points downstream of each deionizing unit. If you took a sample from one of these, what kind of test might you perform on the sampled water?

| Document No.: BioWork-5010 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 3 of 6 |
| --- | --- | --- | --- |
| Reading a Process Diagram |  |  |  |

  - If both deionizers are operating simultaneously, which sampling point could you use to monitor the overall operation?

  - Only a single type of pump is used in this system. What type is it?

  - What is the most common type of valve used in this system?

  - Find the valves that control the reverse osmosis unit being used. What position must these valves be in (open or closed) to make the water flow through the RO unit 5201?

| Document No.: BioWork-5010 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 4 of 6 |
| --- | --- | --- | --- |
| Reading a Process Diagram |  |  |  |

#### Figures/Diagrams and Attachments

  - Lead Sheet (Legend)

  - Water Purification Process diagram

#### Notes/Observations

See text boxes above

N/A

#### Attachments

| Document No.: BioWork-5010 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 5 of 6 |
| --- | --- | --- | --- |
| Reading a Process Diagram |  |  |  |

Figure 9.1 Lead Sheet (Legend)

| Document No.: BioWork-5010 | Revision No.: 00 | Effective Date: 26 Aug 2019 | Page 6 of 6 |
| --- | --- | --- | --- |
| Reading a Process Diagram |  |  |  |

Diagram 9.2	Water Purification Process diagram

| Document No.: BioWork-5030 | Revision No.: 01 | Effective Date: 26 Aug 2019 | Page 1 of 4 |
| --- | --- | --- | --- |
| Conductivity Measurement |  |  |  |

Laboratory / Process Description:

This experiment will introduce conductivity measurement of various types of water. Conductivity is a measure of water's capability to pass electrical flow. This ability is directly related to the concentration of ions in the water. The higher the quality of water, the fewer ions are present. This experiment will allow students to compare the conductivity of different types of water using a conductivity meter.

Signature Log

| Name (print) | Signature | Initials | Date |
| --- | --- | --- | --- |

Bill of Materials and Equipment

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Water for Injection (WFI) or De-ionized | 100 | mL |
| Water (tap) | 100 | mL |
| Normal Saline Solution (0.9 % salt /water) | 100 | mL |
| Beaker (50 - 100 mL) or 50 mL free standing conical tubes | 3 | each |
| Conductivity Meter (Mettler Toledo or equivalent) | 1 | each |

Definitions

Electrical current: is the movement of charged particles (measured in Amps, A)

Conductance: is a measure of how easily those charge particles move through a solution/material (measured in Siemens, S).

| Document No.: BioWork-5030 | Revision No.: 01 | Effective Date: 26 Aug 2019 | Page 2 of 4 |
| --- | --- | --- | --- |
| Conductivity Measurement |  |  |  |

Conductivity: is the conductance (S) measured across a specified distance through a material/solution (measured in Siemens per centimetre, S/cm)

Room Conditions (OPTIONAL)

| Room Conditions Check: Room |  |  | Performed By / Date | Verified By / Date |
| --- | --- | --- | --- | --- |
| Temp: ° C | RH: % |  |  |  |

Procedure

Conductivity Measurement

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Turn on conductivity meter if not on already. Check for instrument calibration if required. |  |  |
| 2. | Pour approximately 25 ml of the solution to be tested into a 50 mL conical tube or beaker. Pour the enough solution to ensure the conductivity probe can be sufficiently immersed in the test solution. Record the solutions below: Solution: Approximate mL: Solution: Approximate mL: Solution: Approximate mL: |  |  |

| Document No.: BioWork-5030 | Revision No.: 01 | Effective Date: 26 Aug 2019 | Page 3 of 4 |
| --- | --- | --- | --- |
| Conductivity Measurement |  |  |  |

| 3. | Remove conductivity probe from it’s storage beaker and rinse the probe with WFI or Lab De-ionized water. Lightly tap dry the probe with a KimWipe or equivalent tissue. |  |  |
| --- | --- | --- | --- |
| 4. | Place the conductivity probe into the test solution and read the conductivity (μS/cm). For each solution tested record the conductivity in the results section. |  |  |
| 5. | Rinse the conductivity probe with de-ionized water and repeat the drying process between each test. |  |  |

Equipment Printouts and/or Attachments:

N/A

Results / Notes / Observations:

Conductivity Measurements

| Reading # | Solution | Conductivity (μS/cm) |
| --- | --- | --- |
| 1. | Water For Injection (WFI) | (μS/cm) |
| 2. | De-Ionized Water | (μS/cm) |
| 3. | Tap Water | (μS/cm) |
| 4. | Normal Saline | (μS/cm) |
| 5. |  | (μS/cm) |
| 6. |  | (μS/cm) |
| 7. |  | (μS/cm) |

| Document No.: BioWork-5030 | Revision No.: 01 | Effective Date: 26 Aug 2019 | Page 4 of 4 |
| --- | --- | --- | --- |
| Conductivity Measurement |  |  |  |

Conclusions:

| Document No.: BRP-BioWork-5040 | Revision No.: 01 | Effective Date: 16 Nov 2018 | Page 1 of 8 |
| --- | --- | --- | --- |
| Mixing and Solubility |  |  |  |

Laboratory / Process Description:

This experiment will introduce mixing, solubility, and the effects of time, temperature, and mixing speed on the dissolution of a solute into a solvent. Each team should chose a different solute from the table below. The solvent in all cases will be water. This experiment will allow students to compare the solubility of different household chemicals in water as well as the mixing parameters.

Signature Log

| Name (print) | Signature | Initials | Date |
| --- | --- | --- | --- |

Bill of Materials and Equipment (Per Experiment)

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Water (Solvent) 3 x 100mL | 300 | mL |
| Common Household Chemical (see solubility Table) | 50-200 | grams |
| Beaker (250 mL) | 3 | each |
| Magnetic Stir plate with temperature control, | 1 | each |
| Magnetic Stir Bars | 3 | each |
| Scale | 1 | each |
| Stopwatch | 1 | each |
| Thermometer (digital) | 1 | each |
| pH/Conductivity Meter (optional for salt solutions) | 1 | each |
| Refractometer (Optional, for sucrose solutions) | 1 | each |

| Document No.: BRP-BioWork-5040 | Revision No.: 01 | Effective Date: 16 Nov 2018 | Page 2 of 8 |
| --- | --- | --- | --- |
| Mixing and Solubility |  |  |  |

Definitions

Solutions: are homogeneous mixtures made up of one or more solutes dissolved in a solvent.

Solubility: is defined to be the maximum amount of solute that will dissolve in a given amount of solvent at a specific temperature. Typically, solubility is recorded as grams of solute per 100g of solvent.

Refractometer: an instrument for measuring refractive indices (as for identification or the determination of sugar content). Max read on the refractometer is 18% w/w.

Solubility Table

| Item | Solubility (g/100mL) |  |  |  |
| --- | --- | --- | --- | --- |
|  | 0°C | 20°C | 70% Max Solubility Wt. (at 20°C) | 100°C |
| Epsom Salt | 27g/100mL | 35g/100mL | 24g/100 mL | 50g/100mL |
| Sugar | 179g/100mL | 204g/100mL | 143/100 mL | 487g/100mL |
| Table Salt | 33g/100mL | 36g/100mL | 25g/100 mL | 40g/100mL |
| Baking Soda | 7g/100mL | 10g/100mL | 7g/100 mL | 17g/100mL |

| Document No.: BRP-BioWork-5040 | Revision No.: 01 | Effective Date: 16 Nov 2018 | Page 3 of 8 |
| --- | --- | --- | --- |
| Mixing and Solubility |  |  |  |

Room Conditions (OPTIONAL)

| Room Conditions Check: Room |  |  | Performed By / Date | Verified By / Date |
| --- | --- | --- | --- | --- |
| Temp: °C | RH: % | ∆P: “WC |  |  |

Procedure

Mixing Study (Mix speed)

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Chose a solute from the table above. Record the solute chosen below: Solute: |  |  |
| 2. | Weight out 70% of the maximum solubility amount for a temperature of 20°C for the solute chosen (ex. Table salt max is 36g; 70% = 25g). Record weight of solute. Solute: Weight: g |  |  |
| 3. | Obtain one of the three (250mL) beakers and pour 100mL of the solvent (water) into it. Place an appropriately sized stir bar into the beaker. Using a digital thermometer measure the initial temperature of the solvent and record below. Starting Temperature: °C |  |  |
| 4. | Place the beaker on a stir plate with the heat OFF. (Room Temperature) |  |  |
| 5. | Set the initial mixing speed to 50 RPM or Low speed RPM: or Speed Setting |  |  |
| 6. | Pour the solute into the beaker and begin timing on the stopwatch. Stop the timer when no visible solute remains and record the results section. |  |  |
| 7. | Repeat the process for mixing speeds of 100 (Medium) and 250 (Hi) RPM using the remaining beakers. Record results in the results section. |  |  |

| Document No.: BRP-BioWork-5040 | Revision No.: 01 | Effective Date: 16 Nov 2018 | Page 4 of 8 |
| --- | --- | --- | --- |
| Mixing and Solubility |  |  |  |

Mixing Study (Mix Temperature)

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Chose a solute from the table above. Record the solute chosen below: Solute: |  |  |
| 2. | Weight out 70% of the maximum solubility amount for a temperature of 20°C for the solute chosen (ex. Table salt max is 36g; 70% = 25g). Record weight of solute. Solute: Weight: g |  |  |
| 3. | Obtain one of the three (250mL) beakers and pour 100mL of the solvent (water) into it. Place an appropriately sized stir bar into the beaker. Using a digital thermometer measure the initial temperature of the solvent and record below. Note: If possible, start with chilled solvent (ie. below room temp) Starting Temperature: °C |  |  |
| 4. | Place the beaker on a stir plate with the heat OFF. (Room Temperature or below) |  |  |
| 5. | Set the mixing speed to 100 RPM or Medium speed RPM: or Speed Setting |  |  |
| 6. | Pour the solute into the beaker and begin timing on the stopwatch. Stop the timer when no visible solute remains and record the results section. |  |  |
| 7. | Repeat the process for mixing temperatures of 30°C and 60°C using the remaining beakers. Keep the RPM constant at 100 RPM (Medium). Record results in the results section. |  |  |

| Document No.: BRP-BioWork-5040 | Revision No.: 01 | Effective Date: 16 Nov 2018 | Page 5 of 8 |
| --- | --- | --- | --- |
| Mixing and Solubility |  |  |  |

Optional Mixing Study (Analytical Measurement)

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Choose either the salt or sucrose (sugar) solute from the table above. Record the solute chosen below: Solute: |  |  |
| 2. | Weight out the solute according to the amounts listed below (based on experimental temperature of 20°C -25°C). Record weight of solute. Salt: NMT 1g/100mL Water Sucrose (Sugar): NMT 18g/100mL Solute: Weight: g |  |  |
| 3. | Obtain an additional beaker and pour 100mL of the solvent (water) into it. Place an appropriately sized stir bar into the beaker. Using a digital thermometer measure the initial temperature of the solvent and record below. Starting Temperature: °C |  |  |
| 4. | Place the beaker on a stir plate with the heat OFF. (Room Temperature 20-25°C) |  |  |
| 5. | Set the mixing speed to 50 RPM (Low Speed) RPM: |  |  |
| 6. | Pour the solute into the beaker and begin timing on the stopwatch. Record the conductivity or concentration using the appropriate instrument at timed intervals during the mixing. Stop the timer when no visible solute remains. Record the time and conductivity / concentration in the results section. Note: For Salt use the Conductivity Meter Note: For Sugar use the Refractometer Instrument Used: |  |  |
| 7. | Repeat the process for the other solute if time permits. |  |  |

Note: The samples may have to be diluted with DI water to be within the readable range of the instrument used.

| Document No.: BRP-BioWork-5040 | Revision No.: 01 | Effective Date: 16 Nov 2018 | Page 6 of 8 |
| --- | --- | --- | --- |
| Mixing and Solubility |  |  |  |

Equipment Printouts and/or Attachments:

N/A

Results / Notes / Observations:

Experiment 1 (Mix Speed)

| Solute: |  |  |  |
| --- | --- | --- | --- |
| Solvent: Water | mL |  | Initial Temp: °C |
| Mix Speed | Temp (°C) | Time to Dissolve (minutes) | Visible solute? Y/N |
| Run 1: RPM |  |  |  |
| Run 2: RPM |  |  |  |
| Run 3: RPM |  |  |  |

Experiment 2 (Mix Temperature)

| Solute: |  |  | Mix Speed: 100 RPM |
| --- | --- | --- | --- |
| Solvent: Water | mL |  | Initial Temp: °C |
| Mix Temperature | Mix Speed (fixed) | Time to Dissolve (minutes) | Visible solute? Y/N |
| Run 1: (°C) |  |  |  |
| Run 2: (°C) |  |  |  |
| Run 3: (°C) |  |  |  |

| Document No.: BRP-BioWork-5040 | Revision No.: 01 | Effective Date: 16 Nov 2018 | Page 7 of 8 |
| --- | --- | --- | --- |
| Mixing and Solubility |  |  |  |

Optional Experiment (Analytical Measurement)

| Solute: SALT |  |  | Mix Speed: 50 RPM |
| --- | --- | --- | --- |
| Solvent: Water | mL |  | Initial Temp: °C |
| Mix Intervals | Time (Minutes) | Conductivity Reading (μS/cm) | Visible solute? Y/N |
| Step 1 |  |  |  |
| Step 2 |  |  |  |
| Step 3 |  |  |  |
| Step 4 |  |  |  |
| Step 5 |  |  |  |
| Step 6 |  |  |  |
| Step 7 |  |  |  |
| Step 8 |  |  |  |
| Step 9 |  |  |  |
| Step 10 |  |  |  |

| Document No.: BRP-BioWork-5040 | Revision No.: 01 | Effective Date: 16 Nov 2018 | Page 8 of 8 |
| --- | --- | --- | --- |
| Mixing and Solubility |  |  |  |

Optional Experiment (Analytical Measurement)

| Solute: SUGAR |  |  | Mix Speed: 50 RPM |
| --- | --- | --- | --- |
| Solvent: Water | mL |  | Initial Temp: °C |
| Mix Intervals | Time (Minutes) | Refractometer Reading (RI) | Visible solute? Y/N |
| Step 1 |  |  |  |
| Step 2 |  |  |  |
| Step 3 |  |  |  |
| Step 4 |  |  |  |
| Step 5 |  |  |  |
| Step 6 |  |  |  |
| Step 7 |  |  |  |
| Step 8 |  |  |  |
| Step 9 |  |  |  |
| Step 10 |  |  |  |

| Document No.: BioWork-5060 | Revision No.: 02 | Effective Date: 25 Jul 2019 | Page 1 of 10 |
| --- | --- | --- | --- |
| Plate and Frame Filtration |  |  |  |

Laboratory / Process Description:

This experiment will introduce mixing, solubility, and tangential flow filtration using a plate & frame type filter system that is commonly used in the pharmaceutical industry. The system used in class is a type used in the beverage industry for clarifying wines. This experiment will allow students to observe the filtration of a suspension in water as well as record mixing parameters.

Signature Log

| Name (print) | Signature | Initials | Date |
| --- | --- | --- | --- |

Bill of Materials and Equipment (Per Experiment)

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Plate and Frame Filtration Kit, Buon Vino MiniJet Wine Filter Model MJT9308 - Filtration Pads (475 cm2 total surface area, 3 pads), Pad set #1-3; set of three | 1 | each |
| Water (Solvent) | 4-5 | L |
| Activated Charcoal, or equivalent (Granular Activated Carbon (GAC) - irregular shaped particles with sizes ranging from 0.2 to 5 mm). | 1-10 | grams |
| Raw Vinegar (Alternative to Activated Charcoal) | 1 | gal |
| Beaker (1000 mL) | 2 | each |
| Glass tray (9” x11”) or equivalent (to pre-soak filter packs) | 1 | each |
| Magnetic Stir plate | 1 | each |
| Magnetic Stir Bars | 1 | each |

| Document No.: BioWork-5060 | Revision No.: 02 | Effective Date: 25 Jul 2019 | Page 2 of 10 |
| --- | --- | --- | --- |
| Plate and Frame Filtration |  |  |  |

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Scale | 1 | each |
| Stopwatch | 1 | each |

Definitions

Suspensions: are heterogeneous mixtures in which solute-like particles settle out of a solvent-like phase sometime after their introduction. Generally a mixture in which small particles of a substance are dispersed throughout a gas or liquid. If a suspension is left undisturbed, the particles are likely to settle to the bottom. The particles in a suspension are larger than those in either a colloid or a solution. Muddy water is an example of a suspension.

Solutions: are homogeneous mixtures made up of one or more solutes dissolved in a solvent.

Solubility: is defined to be the maximum amount of solute that will dissolve in a given amount of solvent at a specific temperature. Typically, solubility is recorded as grams of solute per 100g of solvent.

Activated carbon: Derived from selected grades of coconut shell, it is ideal for many liquid phase applications including the removal of organics from water streams. Certified by the NSF, it is suitable for drinking water and food grade applications. Granular Activated Carbon (GAC) - irregular shaped particles with sizes ranging from 0.2 to 5 mm.

Filter: A device for separating solids from liquids, or particles from gases consisting of a porous substance through which only the liquid or gas can pass through.

Raw Vinegar: Organic, unfiltered apple cider vinegar; also contains "mother," strands of proteins, enzymes and friendly bacteria that give the product a murky appearance; NOT pasteurized or filtered.

Clarification: To remove unwanted solid particles from fluids or to separate a required solid from a fluid. Clarification can be achieved using either filtration or centrifugation techniques.

Filter Pad Table

Filtration Pads (475 cm2 total surface area, 3 pads)

| Filter Pad | Material | Pore Size |
| --- | --- | --- |
| Pad Set #1 | Course | 5.0 micron |
| Pad Set #2 | Fine | 1.8 micron |
| Pad Set #3 | Very Fine | 0.5 micron |

| Document No.: BioWork-5060 | Revision No.: 02 | Effective Date: 25 Jul 2019 | Page 3 of 10 |
| --- | --- | --- | --- |
| Plate and Frame Filtration |  |  |  |

Room Conditions (OPTIONAL)

| Room Conditions Check: Room |  |  | Performed By / Date | Verified By / Date |
| --- | --- | --- | --- | --- |
| Temp: °C | RH: % | ∆P: “WC |  |  |

Procedure

Formulate Activated Charcoal Suspension

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Weigh approximately 1-10 g of activated charcoal and place into a 1 L glass beaker. Record the weight below: Weight of charcoal: g |  |  |
| 2. | Add 950 mL of lab grade water or equivalent to the beaker with the charcoal. Record the added volume below. Volume of water: mL |  |  |
| 3. | Place an appropriately sized stir bar into the beaker. Place beaker on magnetic stir plate and set the mix speed to 25 -50 RPM. Record mixing speed below. Using a digital thermometer measure the initial temperature of the solution and record below. Mixing Speed: RPM Starting Temperature: °C |  |  |
| 4. | Continuously Stir the mixture throughout the filtration experiment |  |  |

| Document No.: BioWork-5060 | Revision No.: 02 | Effective Date: 25 Jul 2019 | Page 4 of 10 |
| --- | --- | --- | --- |
| Plate and Frame Filtration |  |  |  |

Raw Vinegar Suspension

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Obtain at least 1 liter (preferably 1 gal) of Raw Vinegar (with “mother”) and pour into a 1 L glass beaker. If using gallon container can filter directly. Record the volume below: Volume of raw vinegar: Note: If using the gallon container, the vinegar can filter directly from the container. |  |  |
| 2. | Place an appropriately sized stir bar into the beaker or container. Place beaker or container on magnetic stir plate and set the mix speed to 25 -50 RPM. Record mixing speed below. Using a digital thermometer measure the initial temperature of the solution and record below. Mixing Speed: RPM Starting Temperature: °C |  |  |
| 3. | Continuously Stir the mixture throughout the filtration experiment |  |  |
| 4. | Follow Filtration set-up and filtration process instructions below to complete the filtration. |  |  |

| Document No.: BioWork-5060 | Revision No.: 02 | Effective Date: 25 Jul 2019 | Page 5 of 10 |
| --- | --- | --- | --- |
| Plate and Frame Filtration |  |  |  |

Set Up and Operation of Filtration System

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Chose an appropriate filter pad set from the Filter Pad table above and record filter number and pore size below. Filter Pad set: # Pore Size: micron |  |  |
| 2. | Remove the 3 pads from the package and place them in a clean glass tray with enough fresh water to cover the pads and allow them to soak for approximately 1minute. Make sure the pads are not bunched together while soaking, allow the water to saturate on both sides of the 3 pads. All 3 pads in a package are the same grade. Note: Grades cannot be interchanged i.e., use all #1 pads, or all #2 pads, or all #3 pads. |  |  |
| 3. | Attach the small diameter piece of hose supplied to the fitting located under the drip tray (A on diagram). This same hose may be used for cleaning the pump. |  |  |
| 4. | Loosen the two black hand wheels and remove the two central plates. |  |  |

| Document No.: BioWork-5060 | Revision No.: 02 | Effective Date: 25 Jul 2019 | Page 6 of 10 |
| --- | --- | --- | --- |
| Plate and Frame Filtration |  |  |  |

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 5. | Remove the pads from the water and begin inserting them into the filter body starting from the back. The coarse side must be facing the hand wheels and holes will be aligned (see filter insertion diagram). The pads have a tight fit. Make sure to press down on the pad so that both sides of the “T” are sitting on the two side bolts and the pad is straight along the top. Place the pad into the filter body, next place one of the central filtering plates, insert the second pad and the next plate in the same manner, followed by the last pad. Once the pads and plates are inserted, all the PADS and the PLATES should be even at the top and should have the holes as follows: TOP LEFT - BOTTOM RIGHT. NOTE: The pads are cut in a “T” formation with one side of the “T” higher than the other. The hooks on both sides of the filter plates also have the same offset. This allows you to insert the pads and plates properly. (See Filter Pad Insertion Diagram below.) |  |  |
| 6. | Tighten the hand wheels to ensure a tight seal on the pads. |  |  |
| 7. | Insert the intake hose, which contains a stainless steel wire and an anti-sediment tip into the inlet on the pump (“B” on the System diagram below). |  |  |
| 8. | Insert the long, slightly larger diameter outtake tube into the exit nozzle (“C” on the System diagram). |  |  |
| 9. | Perform a water leak test by placing the intake hose into a beaker with 1-2L of DI water and running the filtration system. Tighten the hand wheels to ensure a tight seal on the pads. Remove intake from the DI water once leaks have been sealed. |  |  |
| 10. | Place the intake hose into the bulk solution (Charcoal/Vinegar) to be filtered. Place hose into the beaker without disturbing the magnetic stir bar. The solution must be stirred throughout the filtration process. |  |  |

| Document No.: BioWork-5060 | Revision No.: 02 | Effective Date: 25 Jul 2019 | Page 7 of 10 |
| --- | --- | --- | --- |
| Plate and Frame Filtration |  |  |  |

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 11. | Place the outtake tube and the small tube attached to the drip tray into the clean beaker which will hold the filtered solution (Filtrate). If desired, the drip tray tube may be placed into a separate container and re-filtered at the end. |  |  |
| 12. | Plug in the cord and turn on the switch located at the rear of the filter. Record filtration start time below. Filtration Start Time: Note: Pump speed is fixed at approximately 1-2 L per minute |  |  |
| 13. | If initial pump is running dry, prime the pump if necessary by forcing the filtrate into the pump using the intake hose |  |  |
| 14. | Continue filtration until the bulk of the solution has been filtered. Record the end of filtration time and the temperature of the filtrate below. Filtration Stop Time: Temperature of filtered solution (filtrate): °C |  |  |
| 15. | Record Observations in the results section |  |  |
| 16. | Clean filtration system according to instructions below. |  |  |

| Document No.: BioWork-5060 | Revision No.: 02 | Effective Date: 25 Jul 2019 | Page 8 of 10 |
| --- | --- | --- | --- |
| Plate and Frame Filtration |  |  |  |

Diagram 1: Minijet Filtration System

Diagram 2: Filter Pad Insertion

Correct Insertion	Incorrect Insertion

Course Side	Smooth (Note Holes)

NOTE: Coarse side always faces towards the black hand wheels on all three types of pads.

NOTE: All 3 pads MUST be the same grade. You CANNOT interchange grades.

| Document No.: BioWork-5060 | Revision No.: 02 | Effective Date: 25 Jul 2019 | Page 9 of 10 |
| --- | --- | --- | --- |
| Plate and Frame Filtration |  |  |  |

Clean-up of Filtration System

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Remove the short tube which had been placed on the fitting under the drip tray. Disconnect the short tube which leads from the pump to the first filtering plate and attach the drip tray tube (A in diagram) to the outlet portion of the pump, leaving the intake tube connected to the pump. You now have a direct line through the pump, bypassing the filter plates. |  |  |
| 2. | Pump clean cold water through the pump. Place the intake tube into approx. 15 liters or 3 gallons of clean water and the outtake tube into an empty container. Turn on the pump, allowing the clean water to flush out the pump. |  |  |
| 3. | Turn unit off and unplug power cord. |  |  |
| 4. | Remove the filtering plates and hoses and clean. When cleaning the drip tray be very careful not to wet motor. NOTE: Do not allow any water to get into the motor section. |  |  |

| Document No.: BioWork-5060 | Revision No.: 02 | Effective Date: 25 Jul 2019 | Page 10 of 10 |
| --- | --- | --- | --- |
| Plate and Frame Filtration |  |  |  |

Equipment Printouts and/or Attachments:

Results / Notes / Observations:

N/A

Filtration Experiment

| Activated Charcoal |  | Water: mL Charcoal: g |  |  |
| --- | --- | --- | --- | --- |
| Step | Time | Filtrate Observation (cloudy, color) | Mix Speed (RPM) | Temperature |
| Intitial |  |  |  | °C |
| Middle |  |  |  | °C |
| End |  |  |  | °C |

Resources

Filtration System & Pads:

Buon Vino MiniJet Wine Filter Model MJT9308 Buon Vino Manufacturing Inc. www.buonvino.com

Activated Carbon:

General Carbon Corporation, GC Powdered-S http://www.generalcarbon.com/activated- carbon/gc-powdered- s/?gclid=EAIaIQobChMIlfO6mf2V3wIVAtvACh1etwgEEAAYASAAEgIFvfD_BwE

Chemviron

https://www.chemviron.eu/products/activated-carbon/granulated-activated-carbon/

| Document No.: BioWork-5070 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 1 of 5 |
| --- | --- | --- | --- |
| Column Chromatography Separation (Grape Soda) |  |  |  |

#### Purpose

To Use reverse phase chromatography to separate the blue and red dyes from grape soda (Red #40 and Blue #1)

#### Scope

This procedure applies to the separation of the blue and red dyes which comprise the purple grape color in grape soda.

#### Responsibilities

Technicians will operate equipment under the direction of an instructor. It is the responsibility of students undergoing gowning training to comply with this procedure.

#### Definitions

Chromatography:	Is the collective term for techniques used to separate the components of mixtures. Chromatography is a physical method of separation in which the components to be separated are distributed between two phases, one of which is solid stationary phase, while the liquid mobile phase or the elution phase moves in a definite direction.

Mobile phase:	Transports the solutes or components to be separated through a column or plate of stationary phase material.

#### References

Sep-Pak C18 Cartridge Video by Flinn Scientific - https://youtu.be/JqDaCkH8oSI Red Dye 40 (Allura Red AC, or FD&C Red 40)

- Molar mass: 496.42 g/mol

- Formula: C18H14N2Na2O8S2

- IUPAC ID: disodium 6-hydroxy-5-((2-methoxy-5-methyl-4-sulfophenyl)azo)-2- naphthalenesulfonate

Blue Dye 1 (Brilliant Blue FCF, or FD&C Blue 1)

- Formula: C37H34N2Na2O9S3

- Molar mass: 792.85 g/mol

- IUPAC ID: disodium;2-[[4-[ethyl-[(3-sulfonatophenyl)methyl]amino]phenyl]-[4- [ethyl-[(3-sulfonatophenyl)methyl]azaniumylidene]cyclohexa-2,5-dien-1- ylidene]methyl]benzenesulfonate

| Document No.: BioWork-5070 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 2 of 5 |
| --- | --- | --- | --- |
| Column Chromatography Separation (Grape Soda) |  |  |  |

#### Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| 10 mL Grape soda (NuGrape or equivalent) | 1 | Each |
| 15 mL 10% v/v ethanol | 1 | Each |
| 20 mL 90% v/v ethanol | 1 | Each |
| 20 mL Deionized or Distilled water | 1 | Each |
| 10-mL Plastic Disposable syringe | 1 | Each |
| Test tubes | 2 | Each |
| Test tube rack | 1 | Each |
| Pen and tape for labeling | 1 | Each |
| Sep-Pak Plus, C-18 cartridge (Waters Associates) (Cat. No. WAT 020515) | 1 | Each |
| 100 mL beaker for waste water and organic solvent elutes from the cartridges | 1 | Each |
| 50-mL beakers |  |  |
| Sink for disposal of the eluates; appropriate waste container |  |  |

- Safety

Wear lab PPE including lab coat, gloves, and safety glasses.

| Document No.: BioWork-5070 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 3 of 5 |
| --- | --- | --- | --- |
| Column Chromatography Separation (Grape Soda) |  |  |  |

#### Procedure

  - General Notes.

    - In general, we separate in order to:

      - Identify – What is present in the sample

      - Quantify – How much is present in the sample

      - Purify – Isolate a compound from the sample

    - Key to separation is the fact that there are differing affinities of analyte(s) in a sample for the solid stationary cartridge phase versus the mobile phases. This will allow elution (removal) of the dyes from the cartridge one at a time.

    - Reverse phase mode of chromatography refers to the fact that one uses columns in which non-polar molecules (C18 hydrocarbons bonded to silica particles) have been chemically attached to silica particles packed in the cartridge.

  - Chromatography Procedure

    - Draw 5 mL of 90% ethanol into the syringe using the syringe plunger.(Use volume markings on the syringe barrel.)

    - Attach syringe to the short end of the C-18 cartridge

    - Place a 100-mL waste beaker under syringe/cartridge to catch the flow of ethanol that will run through the cartridge.

    - Gently push the ethanol through the cartridge by using the syringe plunger. This step washes out any impurities bound to the cartridge packing and activates the cartridge packing to put it into a form which will retain the two dyes.

    - Remove the syringe from the cartridge and fill it with 5 mL of DI water and re- attach the syringe to the short end of the cartridge.

    - Push the water through the cartridge as you did before, by using the syringe plunger. Use the 100 ml waste beaker to catch the flow of water from the cartridge. This step ensures that the column is saturated with water and all the 90% ethanol used to activate the cartridge is flushed out.

    - Remove the syringe from the cartridge and fill the syringe with 5 mL of Grape soda by using the plunger to draw it up. After attaching the barrel of the syringe to

| Document No.: BioWork-5070 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 4 of 5 |
| --- | --- | --- | --- |
| Column Chromatography Separation (Grape Soda) |  |  |  |

the cartridge, use the plunger to push the soda through the cartridge into the 100- mL waste beaker.

      - The dyes will remain on the cartridge.

      - The flow from the soda into the beaker should be clear. The flow contains water and dissolved sugar and salts which were in the soda.

    - Remove the syringe from the cartridge and fill it with 7 mL of DI water by using the plunger. After attaching the barrel of the syringe to the cartridge, gently push the water through the cartridge into the 100-mL waste beaker.

    - Now you will remove (elute) the red dye from the cartridge. Remove the syringe from the cartridge and fill it with 7 mL of 10% ethanol using the plunger and after attaching the barrel of the syringe to the cartridge, push the 7 mL of 10% ethanol through the cartridge. In a small test catch what flows through the cartridge.

      - The red dye will be in the solution that comes out of the cartridge.

      - The blue dye will remain in the cartridge.

    - Remove the syringe from the cartridge and fill the syringe barrel with 10 mL of 90% ethanol by using the plunger. After re-attaching the barrel of the syringe to the cartridge, gently push the 90% ethanol solvent through the cartridge. Catch the solvent eluting from the cartridge—(which contain the blue dye)—in a second small test tube.

      - Be prepared to discuss your observations with the class.

Note: The C18 cartridge can be re-used multiple times over a period of months by repeating steps

8.2.1 – 8.2.10.

#### Figures and Diagrams

N/A

| Document No.: BioWork-5070 | Revision No.: 02 | Effective Date: 26 Aug 2019 | Page 5 of 5 |
| --- | --- | --- | --- |
| Column Chromatography Separation (Grape Soda) |  |  |  |

#### Notes/Observations

| Step | Eluent (mobile phase) | Color of Effluent |
| --- | --- | --- |
| 8.2.4 Regeneration/activation | 90% Ethanol |  |
| 8.2.6 DI water Flush | DI Water |  |
| 8.2.7 Loading | Grape Soda |  |
| 8.2.8 DI water Flush | DI Water |  |
| 8.2.9 Red Dye Elution | 10% Ethanol |  |
| 8.2.10 Blue Dye Elution | 90% Ethanol |  |
| Done by: Date: Checked By: Date: |  |  |

- Attachments

N/A

| Document No.: BioWork-6015 | Revision No.: 00 | Effective Date: 18 Oct 2019 | Page 1 of 12 |
| --- | --- | --- | --- |
| Feedback Control (Temperature & Relative Humidity) |  |  |  |

#### Purpose

To demonstrate the use and interaction of a sensor, transmitter, control device and final control element.

Compare automated control with manual control.

#### Scope

In this system, the air humidity in a container is continuously controlled by a fogging system. Air temperature is controlled by a cooling circuit (Fan) and heating circuit by a 40 to 60W bulb. The humidity and cooling temperature is controlled by one controller, while the heating is controlled by a separate rheostat (Manually). In order to maintain the temperature and relative humidity at the set point, a feedback loops control power to a fogger (humidifier) and the cooling fan. A manual rheostat controls the heating bulb. The feedback loop works like this:

- A light bulb provides a heat-source for the container.

- A fogging device provides humidification

- A temperature sensor takes the current temperature of the air in the container

- A device (the transmitter/sensor) relays the information to the controller.

- The controller compares the current temperature and relative humidity to the set point.

- If the temperature rises above the set point, the controller tells another device (actuator) to turn ON power to the cooling fan (final control element); if the relative humidity is below set-point, the fogging unit comes ON.

- Once the temperature falls below the set point, the controller tells the actuator to turn OFF power to the cooling fan; another circuit (manually) will turn ON the light bulb for additional heat.

- Once the relative humidity reaches set point, the controller will turn the fogging unit OFF

#### Responsibilities

It is the responsibility of students undergoing training to comply with this procedure. Technicians will operate equipment under the direction of an instructor.

#### Definitions

Feedback Control System:	A system in which the value of some output quantity is controlled by feeding back the value of the controlled quantity and using it to manipulate an input quantity so as to bring the value of the controlled quantity closer to a desired value.

Relative Humidity:	is the amount of water vapor present in the air divided by the maximum amount that the air could contain at that temperature. Relative humidity is

| Document No.: BioWork-6015 | Revision No.: 00 | Effective Date: 18 Oct 2019 | Page 2 of 12 |
| --- | --- | --- | --- |
| Feedback Control (Temperature & Relative Humidity) |  |  |  |

expressed as a percentage. Relative humidity depends on temperature and the pressure of the system of interest.

#### References

- Mathworks feedback Control

https://www.mathworks.com/videos/understanding-control-systems-part-2-feedback-control- systems- 123501.html

#### Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| 52 Qt Plastic storage bin with lid (22”L x 17”W x 13”D) or Hefty® Hi-Rise™ 32 Qt Storage Container; PP/TPE Latches (17”L x 12”W x 14”D) | 1 | Each |
| Zoo Med (RF-10) Repti Fogger (terrarium Humidifier) | 1 | Each |
| Zoo Med (HT-10) Hygrotherm Humidity & Temperature Controller | 1 | Each |
| Zoo Med (AA-13) Aqua Cool Aquarium Cooling Fan | 1 | Each |
| Zoo Med (RT-10) ReptiTemp Rheostat | 1 | Each |
| REPTI ZOO Reptile Lamp Holder | 1 | Each |
| LCD Digital Thermometer | 1 | Each |

- Safety

Wear lab PPE including lab coat, gloves, and safety glasses when working in the lab..

#### Procedure

  - General Assembly

Note: Alternative assemblies listed in the procedures below are aimed at having the ability to place the lid on (Closed system) without interference.

    - Obtain all items listed in the materials table.

| Document No.: BioWork-6015 | Revision No.: 00 | Effective Date: 18 Oct 2019 | Page 3 of 12 |
| --- | --- | --- | --- |
| Feedback Control (Temperature & Relative Humidity) |  |  |  |

    - Clamp the lamp holder to the top rim of the plastic storage bin. Pull power cord so it is located on the outside of the bin. Alternatively, clamp the fixture to a flat piece of plastic and place the unit in the bottom corner of the storage bin. Make a notch along the top edge of the bin to pull the cord up through.

    - Plug the power cord of the lamp into one of the sockets on the Rheostat. Place a 40 or 60 Watt incandescent light bulb (no LED) into the lamp holder

    - Place the Hygrotherm controller on the outside of the storage bin (the environment). Place the temperature/RH sensor inside the environment. Alternatively, a hole or notch could be made in the top of the bin to allow the probe wire to pass without interfering with the placement of the top.

    - Assemble the Repti-fogger per manufacturer’s instructions. Clamp the outfeed hose to the top rim of the storage bin in such a manner that the hose protrudes into the space 3-4 inches. Alternatively cut a circular hole in the side of the tub and push the hose into the interior. Plug the power cord into the Hygrotherm controller receptacle marked “Humidity”

    - Assemble the cooling fan unit per manufacturer’s instructions and secure output air tube to the rim of the storage bin in a manor so that the output (cooling) air is directed into the environment (storage bin). Plug the power cable into the Hygrotherm Controller Receptacle marked “Temperature”

Note: The cooling assembly is meant to be hung on the outside from the top edge of the storage bin.

    - Proceed to the next section for programming the controller

  - Programming the Hygrotherm Controller: Temperature

Note: The Thermostat function of your HYGROTHERM™ will turn heating or cooling devices on and off to maintain the optimum temperature

    - The temperature control range is 50°F to 122°F (10°c to 50°C).

    - Press and release the orange "Mode" button to cycle between the three main modes (°C, %RH, °F). Each time the "Mode" button is pressed, the next setting will appear. Continue pressing the "Mode" button until either 0 c or °F shows next to the larger set of numbers on the left side of the display above the word "ACTUAL").

| Document No.: BioWork-6015 | Revision No.: 00 | Effective Date: 18 Oct 2019 | Page 4 of 12 |
| --- | --- | --- | --- |
| Feedback Control (Temperature & Relative Humidity) |  |  |  |

    - To adjust the temperature setting, press the orange "Set" button while in °C or °F mode. When the "Set" button is pressed, the programmed temperature will appear above the word "ACTUAL." Press the "Set" button once to change the programmed temperature one degree. Press and hold the button to increase by several degrees. If you pass the desired setting, press and hold the "Set" button to cycle through the numbers again. The controller will automatically return to the normal display about 3 seconds after the desired temperature is set, and the programmed temperature will appear above the word "PROGRAM."

    - The temperature program is now set and can be changed at any time by pressing the "Set" button while in the appropriate mode.

    - NOTE: While in Fahrenheit (°F) mode, the temperature limit on the LCD above the word "PROGRAM" is 99°F. If the temperature program is set to 100°F or more, the display will show"--" above the word "PROGRAM." To view the programmed temperature setting when this occurs, press and release the "Set" button and the programmed temperature will appear above the word "ACTUAL." The normal display will resume after a few seconds.

  - Programming the Hygrotherm Controller: Humidity

Note: The Hygrostat function of the HYGROTHERM™ will turn humidifiers on and off to maintain the optimum humidity level

    - The humidity control range is 15% to 95% RH (Relative Humidity)

    - Press and release the orange "Mode" button to cycle between the three main modes (°C, %RH, °F). Each time the "Mode" button is pressed, the next setting will appear. Continue pressing the "Mode" button until the letter "H" and %RH appear next to the larger set of numbers on the left side of the display above the word "ACTUAL"

    - To adjust the relative humidity setting, press the orange "Set" button and the programmed humidity setting will appear above the word "ACTUAL." Press the "Set" button once to increase the setting one percent. Press and hold the button to increase by several percentage points. If you pass the desired setting, press and hold the "Set" button to cycle through the numbers again. The controller will automatically return to the normal display about 3 seconds after the relative humidity is set, and the programmed humidity setting will appear above the word "PROGRAM.

    - The humidity program is now set and can be changed at any time by pressing the "Set" button while in the appropriate mode.

| Document No.: BioWork-6015 | Revision No.: 00 | Effective Date: 18 Oct 2019 | Page 5 of 12 |
| --- | --- | --- | --- |
| Feedback Control (Temperature & Relative Humidity) |  |  |  |

  - Additional Programming Features: HEAT/COOL Mode (FO1)

    - Enter the programming mode by press and hold the mode button (about 3 seconds) while in °C or °F mode until F01 appears on the left side of the display.

    - Press and release the mode button to cycle through the four available features until HEAT/COOL Mode is selected.

    - HEAT/COOL Mode: In HEAT mode, the controller will turn heating devices (e.g., heat lamps, under tank heaters, ceramic heat emitters) on when the actual temperature in the environment drops below the programmed temperature setting.

    - In COOL mode, the controller will turn on a cooling device (e.g., fan) when the temperature in the environment rises above the programmed temperature setting.

    - To switch between HEAT and COOL mode, press and hold the "Mode" button until F01 appears, then release the "Mode" button. Press the "Set" button to switch between modes. When COOL mode is selected, the numbers "00" and the word "COOL" will appear on the display. When HEAT mode is selected, the numbers "01" and the word "HEAT" will appear on the display.

  - Alternative Set-Up (Control Heating Mode)

    - The controller could be set for heat mode mode (see Step 8.4) for programming instructions. In this mode the controller is controlling the heating device.

    - Plug the power cable of the lamp fixture (heat) into the Hygrotherm Controller Receptacle marked “Temperature”.

    - The cooling circuit will be manual. Plug the cooling fan power cable into the rheostat.

NOTE: the rheostat MUST be set to full power and cannot be controlled in this situation. Use the power switch on the fan unit to turn the fan on/off manually.

  - Procedure for the Lab

    - Determine the system mode (Control Cooling /RH) or (Control Heating/RH) and record in the notes observations section.

    - Record system set-points programmed into the controller: Temperature, Humidity (RH) in the notes observations section.

| Document No.: BioWork-6015 | Revision No.: 00 | Effective Date: 18 Oct 2019 | Page 6 of 12 |
| --- | --- | --- | --- |
| Feedback Control (Temperature & Relative Humidity) |  |  |  |

    - Start feedback loops and manual circuits by ensuring power is available to all components.

    - Depending on the temperature and humidity chosen, it may take the system several minutes to reach the set points.

    - Allow the control loop to operate through several cycles until a regular pattern is established.

    - Observe the conditions over time, of the environment and component status. Record the observations in the table located in the Notes Observations section.

    - If time permits, create a graph of time vs. one of the controlled variables in Excel. See example in the Attachments section.

    - Answer/Discuss the questions in the Notes Observation section.

#### Figures and Diagrams

Diagram 1: Control loop Diagram Diagram 2: Repti-Fogger Parts Diagram 3: Aqua Cool Cooling Fan Diagram 4: Hygrotherm Controller

#### Notes/Observations

10.1	Observation Worksheet

#### Attachments

11.1	Excel Graph example

| Document No.: BioWork-6015 | Revision No.: 00 | Effective Date: 18 Oct 2019 | Page 7 of 12 |
| --- | --- | --- | --- |
| Feedback Control (Temperature & Relative Humidity) |  |  |  |

TI

Temperature

Indicator

Temperature

Sensor

Humidity

Sensor

Cooling

TC

HC

Fan

Temperature	Humidity

Controller	Controller

Temperature

Sensor

Temperature	(Optional)

Sensor

Heating

Heater

Rheostat

Humidification

Humidifier

(fogger)

Activity: 6015 Feedback Control Loop LAA

Environment

Diagram 1: Control Loop Diagram

| Document No.: BioWork-6015 | Revision No.: 00 | Effective Date: 18 Oct 2019 | Page 8 of 12 |
| --- | --- | --- | --- |
| Feedback Control (Temperature & Relative Humidity) |  |  |  |

Diagram 2: Repti-Fogger Parts

Diagram 3: Aqua Cool Cooling Fan

| Document No.: BioWork-6015 | Revision No.: 00 | Effective Date: 18 Oct 2019 | Page 9 of 12 |
| --- | --- | --- | --- |
| Feedback Control (Temperature & Relative Humidity) |  |  |  |

Diagram 4: Hygrotherm Controller

| *ocument No.: BioWork-6015 | Revision No.: 00 | Effective Date: 18 Oct 2019 | Page 10 of 12 |
| --- | --- | --- | --- |
| Feedback Control (Temperature & Relative Humidity) |  |  |  |

  - Observations Worksheet

Using the Control Loop Diagram for the Temperature/Humidity Control System as a reference, locate and identify the following parts of the system as set up in your classroom:

    - Temperature-Humidity Sensor/Transmitter

    - Controller/Actuator

    - Heating Circuit/control

    - Cooling Circuit/control

    - Final control elements

    - Temperature sensor/Indicator System Data

| MODE: | □ | Control Cooling /RH |  | □Control Heating/RH |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| SET POINTS: |  | Temperature: °C |  |  | Humidity: %RH |  |
| Time (24hr) | Temp (°C) |  | %RH | Fan (on/off) | Lamp (on/off) | Fog (on/off) |

Questions

| What are the Controlled Variables? | Temperature and Relative Humidity |
| --- | --- |
| What are the Measured Variables? | Temperature and Relative Humidity |
| What is the Manipulated Variable? | Heat when controlling cooling circuit, Fan when controlling Heat circuit |
| What is the Final Control Elements? | Fogger, lamp and or Fan |
| How long did it take temperature to reach set point? | Look at Data |

| Notes: |  |
| --- | --- |
|  | BioWork Lab |

| Document No.: BioWork-6015 | Revision No.: 00 | Effective Date: 18 Oct 2019 | Page 11 of 12 |
| --- | --- | --- | --- |
| Feedback Control (Temperature & Relative Humidity) |  |  |  |

| Explain what happened once Temperature set point was reached. | Controlling device turned off (Lamp or Fan depending on Mode) |
| --- | --- |
| How long did it take Humidity to reach set point? | Look at Data |
| Explain what happened once Humidity set point was reached. | Controlling device turned off (Fogger) |
| What might disturb or upset the process? | Heat output of the lamp is greater than the cooling capacity of the fan; RH is effected by temperature, cooling fan may disperse humidity, lid on or off (closed or open system) |
| Using the ideas generated in the above question, disturb the system: - Describe what was done to disturb the system. - Describe what happened | Answers based on specific senarios |

Optional

Provide a sketch of the graph produced from the data that was entered into an Excel spreadsheet. Indicate the set point, as well as the following time points on the graph using the letters provided:

- Light bulb turned ON by instructor.

- Cooling fan turned ON by Controller.

- Cooling fan turned OFF by Controller (first cycle only).

- Fog Unit turned ON by Controller

- Fog Unit turned OFF by Controller

D.	System was intentionally disturbed. See example in Attachments Section.

Sketch of the graph produced from the data, or attach spreadsheet graph

| Document No.: BioWork-6015 | Revision No.: 00 | Effective Date: 18 Oct 2019 | Page 12 of 12 |
| --- | --- | --- | --- |
| Feedback Control (Temperature & Relative Humidity) |  |  |  |

Attachment 11.1 Excel Graph example

| Document No.: BioWork-7005 | Revision No.: 00 | Effective Date: 29 Aug 2019 | Page 1 of 3 |
| --- | --- | --- | --- |
| IEH/Microbiologique Air Sampler System |  |  |  |

#### Purpose

To learn the principles of air monitoring and the operation of the air sampler.

#### Scope

The IEH Air Sampler is designed to collect microorganisms in air so that they can be transported back to the laboratory for analysis. Air sampling plates obtained from IEH are mounted in the sampler using aseptic technique. The sampler is started and runs for a pre- set period of time, after which the plates are incubated and analyzed.

#### Responsibilities

It is the responsibility of students undergoing training to comply with this procedure.

#### Definitions

Active Air Sampler:	a microbial air sampler that is used to force air into, or onto its collection medium (e.g., Petri Dish with nutrient agar based test media) over a specified period of time.

#### References

- Ordering information for the air samplers https://www.weberscientific.com/air-sampler-microbiologique

#### Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Microbiologique Air Sampler System | 1 | Each |
| 100 mm TSA (tryptic Soy Agar) plates | 2 | Each |
| Optional: Gas Sample Adapter | 1 | Each |

- Safety

Wear lab PPE including lab coat, gloves, and safety glasses when working in the lab.

| Document No.: BioWork-7005 | Revision No.: 00 | Effective Date: 29 Aug 2019 | Page 2 of 3 |
| --- | --- | --- | --- |
| IEH/Microbiologique Air Sampler System |  |  |  |

#### Procedure

Preparation of the Air Sampler

  - Make sure the battery is charged. Turn on the switch of the air sampler briefly and observe the LED light on the side. If operating correctly a green light will come on indicating normal operation. When a run has ended the light will turn blue.

  - The air sampler is normally set to run for 5 minutes, during which time it will sample 500 L of air (100 L/min). Run times between 1 and 10 minutes may be selected by adjusting the dial wheel on the side of the sampler using a small screw driver.

  - If the battery is charged and the sample time is set follow these steps to prepare the air sampler:

    - (1) Make sure the switch is in the “off’ position.

    - (2) Pre-label an IEH air sampling plate with a unique field sample identifier and the date of sampling using Sharpie or other indelible marker on the BOTTOM of the plate (side where agar is). If you label the cover then it may get switched during handling and samples may be mis-identified.

  - Record the unique field sample identifier along with all pertinent attribute information (such as the sample location description, time, date of sample, ambient conditions, etc.) in a field notebook or data sheet.

  - Gently remove the perforated head of the air sampler taking care not to break the connection between the head and the motor body:

    - Place the entire sampler on a flat surface.

    - Grasp the bottom of the sampling head with one hand, and the top with the other, and gently pull up on the perforated head and remove it.

    - Sanitize the head by wiping it with an alcohol swab.

  - Place the labeled agar plate onto the platform and remove the IEH air sampling plate lid. Place the lid in a WhirlPak or Ziplok bag to maintain its sterile integrity while the sample is being collected.

  - Place the head back onto the air sampler:

    - Place the entire sampler on a flat surface.

    - Grasp the bottom of the sampling head with one hand, and the top with the other.

    - With care, place the top of the sampling head over the fresh IEH air sampling plate, taking care not to touch the surface, and push it onto the air sampler. It should be ‘finger tight’, but do not over force.

    - The IEH air sampling plate should be flat and there should be a visible and level gap of about 1/8” between the plate and the holed disk on the air intake.

Sample Collection

| Document No.: BioWork-7005 | Revision No.: 00 | Effective Date: 29 Aug 2019 | Page 3 of 3 |
| --- | --- | --- | --- |
| IEH/Microbiologique Air Sampler System |  |  |  |

  - Place the sampler in the desired location and turn the sampler on. The LED light should show green to indicate that the sampler is running.

  - When the sampler stops the LED light will show blue. Turn off the unit, carefully unscrew the clear housing/sampling port from the ABLE body and replace the plate lid onto the plate. Remove the complete plate from the sampler and load a new plate if another sample is to be collected, otherwise, replace the housing.

Note: As a check on appropriate sampling air flow and sampling head seal, make sure that air dimple marks are clearly visible on the agar surface after sample collection. The pattern should match the holes in the sampling head.

  - The head and housing may be disinfected between sampling events by wiping down with 70% alcohol and allowing to air dry. This is may or may not be necessary between samples in a sampling event.

  - Tape the collected plate shut so the lid cannot become separated from the bottom of the plate. Place the plate into the ziplok or whirlpak bag and secure (one sample plate per bag). Alternatively, after securely taping the lids to the plates, stack the plates 4-5 high and tape them together, and place into a Ziploc or whirlpak bag.

Note: Plates may be kept cool until all collection is complete.

  - Incubate all plates at 37°C for 24-72 hours

  - Observe bacterial/fungal colony formation at the end of the incubation period.

#### Figures and Diagrams

Diagram 1: Air Sampler & Sampler for Gasses

| Document No.: BioWork-7010 | Revision No.: 01 | Effective Date: 29 Aug 2019 | Page 1 of 7 |
| --- | --- | --- | --- |
| Personnel & Environmental Monitoring, Micro Overview |  |  |  |

Laboratory / Process Description:

Environmental Monitoring ensures that a production area is clean and ready for operations. In this laboratory you will perform environmental monitoring in uncontrolled areas, as well as perform personnel monitoring. All plates will be evaluated after appropriate incubation for potential contamination.

Signature Log

| Name (print) | Signature | Initials | Date |
| --- | --- | --- | --- |

Activities List

| Activity | Team #1 | Team #2 | Team #3 | Team #4 |
| --- | --- | --- | --- | --- |
| Personnel Monitoring (Swab & Finger Plate) | x | x | x | x |
| EM : Surface sampling (contact plates) | x | x | x | x |
| EM : Air sampling for non-viable; particulate | x | x | x | x |
| EM : Air sampling for viable | x | x | x | x |
| Micro Overview | x | x | x | x |

| Document No.: BioWork-7010 | Revision No.: 01 | Effective Date: 29 Aug 2019 | Page 2 of 7 |
| --- | --- | --- | --- |
| Personnel & Environmental Monitoring, Micro Overview |  |  |  |

Bill of Materials and Equipment List

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| TSB Agar Plates (Finger Plates) | 2 | Each |
| RODAC Plates (Contact Plates) | 2 | Each |
| Sterile Swabs | 2 | Each |
| Active Air Sampler Plates (M Air T, SAS100) (Optional) | 2 | Each |
| Laser Particle Counter (Optional) | 1 | Unit |
| Active Air Sampler (Optional) | 1 | Unit |

Procedure

Swabbing Technique (Personnel Monitoring)

| Step | Procedure / Process Description | Date | Performed By: | Verified by: |
| --- | --- | --- | --- | --- |
| 1. | Moisten a cotton or fabric tipped sterile swab by dipping the swab into a tube of sterile TSB (Tryptic Soy Broth). |  |  |  |
| 2. | Swab a selected body part, such as nose, ear, mouth, or any skin surface. Record body Part: |  |  |  |
| 3. | Rotate the swab while swabbing body surface so the sample is deposited around the circumference of the swab |  |  |  |
| 4. | Streak the surface of a TSA (Tryptic Soy Agar) plate in a zigzag pattern across the plate. Rotate the swab while streaking to transfer the contents from the entire circumference of the swab. |  |  |  |
| 5. | Label the “agar” side of the plate with the following information: - Team number - Your name - Area of body swab, for example: “outer ear” - Date |  |  |  |

| Document No.: BioWork-7010 | Revision No.: 01 | Effective Date: 29 Aug 2019 | Page 3 of 7 |
| --- | --- | --- | --- |
| Personnel & Environmental Monitoring, Micro Overview |  |  |  |

| 6. | Place the swab plates with swab plates of other team members for incubation. (Incubate upside down at 37 °C) |  |  |  |
| --- | --- | --- | --- | --- |

Finger Plates (Personnel Monitoring)

- Finger Plate Diagram

| Step | Procedure/Process Description | Date | Perform by: | Verify by: |
| --- | --- | --- | --- | --- |
| 1. | Obtain TSB plate and Label bottom of the plate with name and initial |  |  |  |
| 2. | Dab unprotected hands onto surface. Place (4 fingers first, then sweep thumb across surface) (See Finger Plate Diagram above) |  |  |  |
| 3. | Cover plate and incubate upside down at 32 to 37 °C |  |  |  |
| 4. | Record results after 24 to 48 hr incubation Record total no. of colonies seen: CFU Record the no. of different types of colonies: # Types NOTE: If the number of colonies are too numerous to count, record “TNTC” |  |  |  |

| Document No.: BioWork-7010 | Revision No.: 01 | Effective Date: 29 Aug 2019 | Page 4 of 7 |
| --- | --- | --- | --- |
| Personnel & Environmental Monitoring, Micro Overview |  |  |  |

Environmental Monitoring: AIR SAMPLING STATIC CONDITIONS Non-viable particulate (Laser) and simultaneous viable (Active Air Sampler)

| Step | Procedure/Process Description |  | Date | Perform By: | Verify by: |
| --- | --- | --- | --- | --- | --- |
| NOTE : Sample for non-viable and viable particulates at the same time |  |  |  |  |  |
| 1. | Using a Laser Particle counter, sample an area of choice in the Laboratory/Facility or any uncontrolled area outside the Lab. |  |  |  |  |
| 2. | Record the area sampled: |  |  |  |  |
|  | Room number /area |  |  |  |  |
|  | Floor |  |  |  |  |
| 3. | Record Results in the table below and attach print-out to the back of Lab Sheet |  |  |  |  |
|  | STATIC RESULTS | Particle Counts (≥ 0.5µ) |  |  |  |
|  | Room Sampling: |  |  |  |  |
|  | Sample Point #1 |  |  |  |  |
|  | Sample Point #2 (optional) |  |  |  |  |
| 4. | Simultaneously, use the Active Air Sampler (M Air T or equivalent) to sample the same area of choice in the Laboratory/Facility. |  |  |  |  |

| 5. | Label Plates “STATIC”; Label should include - Team # - Date - Sample location / activity |  |  |  |
| --- | --- | --- | --- | --- |
| 6. | Accumulate all plates for incubation. Cover plate and incubate upside down at 32 to 37 °C |  |  |  |

Environmental Monitoring: AIR SAMPLING KINETIC CONDITIONS (Activity) Non-viable particulate and simultaneous viable

| Document No.: BioWork-7010 | Revision No.: 01 | Effective Date: 29 Aug 2019 | Page 5 of 7 |
| --- | --- | --- | --- |
| Personnel & Environmental Monitoring, Micro Overview |  |  |  |

| Step | Procedure/Process Description |  | Date | Perform By: | Verify by: |
| --- | --- | --- | --- | --- | --- |
|  | NOTE : Sample for non-viable and viable particulates at the same time |  |  |  |  |
| 1. | In the same area sampled under static conditions, choose an activity that has potential to generate particulates. Record activity below: |  |  |  |  |
|  | Activity description: |  |  |  |  |
| 2. | Using the Laser Particle counter, Sample air near the “active” team member or activity. |  |  |  |  |
| 3. | Record Particle counts in the table below and attach print-out to the back of Lab Sheet. |  |  |  |  |
|  | KINETIC RESULTS | Particle Counts (≥0.5µ) |  |  |  |
|  | Activity |  |  |  |  |
| 4. | Simultaneously, use the Active Air Sampler (M Air T or equivalent), and sample air near the same “active” team member or activity. |  |  |  |  |
| 5. | Label Plates “KINETIC”; Label should include - Team # - Date - Sample location / activity |  |  |  |  |
| 6. | Accumulate all plates for incubation. Cover plate and incubate upside down at 32 to 37 °C |  |  |  |  |

| Document No.: BioWork-7010 | Revision No.: 01 | Effective Date: 29 Aug 2019 | Page 6 of 7 |
| --- | --- | --- | --- |
| Personnel & Environmental Monitoring, Micro Overview |  |  |  |

Surface Monitoring (RODAC Plates)

| Step | Procedure/Process Description | Date | Perform By: | Verify by: |
| --- | --- | --- | --- | --- |
| 1. | Obtain one RODAC plate per person |  |  |  |
| 2. | Go into the common areas of the facility and choose an area or object to sample. Area / Object: |  |  |  |
| 3. | Open the plate and gently touch the agar to the surface. Slightly rock plate to ensure good contact with the surface. Close the plate and label. |  |  |  |
| 4. | Label Plates “SURFACE”; Label should include - Team # - Initials and Date - Surface/Object Sampled |  |  |  |
| 5. | Spray the surface contacted by the media with 70% IPA and wipe clean. |  |  |  |
| 6. | Record results after incubation for minimum of 18 hours at 32 to 37°C. |  |  |  |

| Document No.: BioWork-7010 | Revision No.: 01 | Effective Date: 29 Aug 2019 | Page 7 of 7 |
| --- | --- | --- | --- |
| Personnel & Environmental Monitoring, Micro Overview |  |  |  |

Incubation of all plates

| Step | Procedure/Process Description | Date | Perform By: | Verify by: |
| --- | --- | --- | --- | --- |
| 1. | Incubate all plates overnight in a 32 - 37° C incubator. (Micro Lab, Room 203). Time in: |  |  |  |
| 2. | Record results after incubation for minimum of 18 hours at 32 to 37°C. For all plates record the Total number of colonies seen (CFU) and the number of different colony types (# Types). NOTE: If the number of colonies are too numerous to count, record “TNTC” |  |  |  |

Attachments / Print Outs:

Results & Discussion

| Notes: |  |
| --- | --- |
|  | BioWork Lab |

| Document No.: BioWork-7020 | Revision No.: 01 | Effective Date: 29 Aug 2019 | Page 1 of 3 |
| --- | --- | --- | --- |
| Gowning for the Cleanroom |  |  |  |

1.0	Purpose

  - To learn proper gowning procedures for cleanroom work, you will view a video and practice gowning and degowning.

- 0	Scope

  - This activity has two parts.

    - In part 1 you will watch a video that demonstrates typical procedures for gowning and degowning.

      - Note: The garments shown in the video are for working in a cleanroom with a very high level classification. If you work in a facility that does not require such a high level (perhaps only Class 100,000 or ISO 8), the gowning requirements may not be as stringent. On the other hand, if you are working in a facility that handles biohazards, there may be additional requirements, especially for exiting the cleanroom.

    - In part 2, everyone in the class will have a chance to practice gowning and degowning. Half of the class gown while the other half observes for violations. Then the groups will switch roles.

  - Responsibilities

  - It is the responsibility of students undergoing gowning training to comply with this procedure.

  - Definitions

  - Glo Germ - an odorless lotion or powder which glows brightly when exposed to ultraviolet light.

  - References

  - BioNetwork Gowning for the Cleanroom video - ncbionetwork.org/educational- resources/elearning/videos/aseptic-gowning-cleanroom

  - Materials

  - Hair and beard covers

  - Gloves

  - Mask

  - Hood

  - Coveralls

  - Glo-Germ

  - Shoe covers

  - Safety goggles

  - Safety

  - Persons practicing gowning should consider leaning on a wall or if necessary sit on a bench to maintain balance. In production these options may not be available, however

| Document No.: BioWork-7020 | Revision No.: 01 | Effective Date: 29 Aug 2019 | Page 2 of 3 |
| --- | --- | --- | --- |
| Gowning for the Cleanroom |  |  |  |

it is expected that no garments intended to be used in sterile processes, be allowed to touch the floor.

  - Procedure

  - Watch the video, collect gowning items and Glo-Germ from your instructor, and then follow the steps below.

    - Important: Do not let outer garments (hair and beard covers, gloves, masks, hoods, coveralls) touch the floor.

    - Do not change your shoes or remove cosmetics for this activity. Put the garments over your street clothes. You will not need to wear an under-suit.

    - Practice in an area divided by a bench or chairs into CLEAN and DIRTY zones.

    - Cover your hair with a hair cover, ensuring to tuck all hair inside the cover. Put on a beard cover if you need one.

    - Wash hands then disinfect hands, wrists, and arms to the elbow with the alcohol solution provided.

      - Note: Washing your hands requires special attention and there may be a company SOP for this process.

    - To see how much care is needed to get your hands really clean:

      - First apply a few drops of Glo-Germ lotion to the palm of one hand.

      - Rub your hands together all over, including the backs of your hands and around your fingers.

      - Scratch your skin gently to get some “germs” under your fingernails.

      - Wash your hands and then check all parts of your hands under the UV light. Record results below.

      - Glo-Germ Present?	Yes	No

    - Continue with the exercise and then see whether any Glo-Germ gets onto the outside of your gowning.

    - Put on your first pair of sterile gloves and select the garments that you need.

      - Note: Only handle the inside of all garments. When donning the sterile garments, it is imperative that personnel take precautions against contaminating the exterior of the garments. The outside of garments must remain contaminant-free.

    - Put on hood: Remove the sterile hood from its package by touching the exterior of the hood as little as possible. Place over the head and adjust by tying behind the head.

    - Put on mask: Open the sterile face mask package and, with shiny side out, place the mask on face. Tie the two pair of strings firmly behind the head. Mold the mask to the bridge of the nose.

    - Put on coverall: Remove the sterile coveralls from the package by grasping the inside of the collar. Hold the coveralls at arm’s length and allow it to unfold. Don't allow the exterior of the coveralls to touch the floor. Unzip the coveralls and, touching only the inside of the coveralls, gather up the coveralls on itself. Place one leg, then the other, into the coveralls' legs. Do not sit down to don

| Document No.: BioWork-7020 | Revision No.: 01 | Effective Date: 29 Aug 2019 | Page 3 of 3 |
| --- | --- | --- | --- |
| Gowning for the Cleanroom |  |  |  |

coveralls. Place arms through the sleeves. Tuck the hood into the coveralls and zip up the front of the coveralls completely to the neck.

    - Put on booties: Remove the sterile boots from their package. Sitting on the bench, put on one boot over the shoe covers. Do not allow the boot to touch the non-sterile gowning side of the floor. Pass the leg over to the clean side of the bench. Repeat for the other boot. Tie boots securely.

    - Put on safety glasses/goggles, face shield, or other face protection.

    - Inspect yourself in a mirror to make sure garments have been correctly placed and fit properly. (If there is no mirror, ask a classmate to check you.)

    - Sanitize gloved hands with sterile 70% IPA.

    - Put on the final pair of sterile gloves: Leaving the first pair of gloves on, pick up a sterile glove with the opposite hand by grabbing the interior cuff that has been rolled back. Place the appropriate hand into the glove, leaving the cuff in the folded position. Repeat the procedure with the glove for the other hand. Pull the sleeves of the coveralls down as far as possible. Roll (unfold) the glove cuffs back ensuring that the sleeves of the gown are tucked inside the gloves. Pull the gloves up over the cuffs of your suit.

    - Disinfect gloves.

    - Optional: have a class member use a UV lamp to check the outside of your gowning for traces of Glo-Germ. Record results

      - Glo-Germ Present?	Yes	No

      - Areas of Glo-Germ Contamination:

    - Stay in your sterile gowning until your instructor tells you to degown so you can get a feel for what it is like to work in these garments.

  - Attachments

9.1	N/A

| Document No.: BioWork-7030 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 1 of 5 |
| --- | --- | --- | --- |
| Aseptic Transfers on an Open Bench |  |  |  |

#### Purpose

To learn proper aseptic technique for transferring sterile liquids or cultures from one container to another—one of the most common tasks you will perform.

#### Scope

In this activity, you will practice aseptic transfer techniques for pipetting and pouring sterile growth media on an open bench. (The term open bench means you are working outside of a laminar flow hood or cleanroom.) After the transferred media has incubated, you will evaluate your aseptic transfer technique by observing the media for contamination. The activity has five parts:

- Part 1: Precautions for Aseptic Transfer

- Part 2: Instructor Demonstration and Student Practice

- Part 3: Aseptic Transfer by Pipetting

- Part 4: Aseptic Transfer by Pouring

- Part 5: Evaluating Your Aseptic Technique

#### Responsibilities

Technicians will operate equipment and conduct the exercise under the direction of an instructor. It is the responsibility of students undergoing training to comply with this procedure.

#### Definitions

Aseptic:	Free from contamination with microorganisms. Asepsis or aseptic means free from pathogenic microorganisms.

Sterile:	Free from bacteria or other microorganisms Sterility can be defined as the freedom from the presence of viable microorganisms. Sterilization is the process of eliminating living microorganisms of all shapes and forms.

Turbidity: The quality of being cloudy, opaque, or thick with suspended mater.

#### References

- “Guideline on Sterile Drug Products Produced by Aseptic Processing,” Center for Drugs and Biologics and Office of Regulatory Affairs, Food and Drug Administration, Sept. 2004.

https://www.fda.gov/regulatory-information/search-fda-guidance-documents/sterile-drug- products- produced-aseptic-processing-current-good-manufacturing-practice

| Document No.: BioWork-7030 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 2 of 5 |
| --- | --- | --- | --- |
| Aseptic Transfers on an Open Bench |  |  |  |

#### Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| 50-mL bottle of tap water | 1 | Each |
| 50-mL bottles of sterile Tryptic Soy Broth growth medium | 2 | Each |
| Sterile 10-mL pipets | 3 | Each |
| Automatic pipet aid (share with one other student) | 1 | Each |
| Sterile 15-mL test tubes (For Pipetting) | 8 | Each |
| Sterile 50-mL Bottles, or 50 mL pre-sterilized (plastic) tubes, (For Pours) | 5 | Each |
| Test tube rack | 2 | Each |
| Marking pen | 1 | Each |
| Labeling tape | 1 | Roll |
| Latex or nitrile gloves | 1 | Pair |
| Spray can of foaming alcohol or spray bottle of 70% Isopropanol (share with one other student) | 1 | Each |

    - Safety

Through the entire activity, you must wear safety glasses, completely fastened lab coats, and latex gloves (nitrile if you are allergic to latex).

#### Procedure

      - Part 1 – Precautions for Aseptic Transfer (General Instruction)

        - Work quickly to minimize the time that containers are open.

        - Hold containers on a slant, not straight up and down, so that there is less chance of air particles falling into the containers.

        - To suction liquid into a pipet, always use a pipet aid—never your mouth.

| Document No.: BioWork-7030 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 3 of 5 |
| --- | --- | --- | --- |
| Aseptic Transfers on an Open Bench |  |  |  |

        - Do not touch the mouths of containers or insides of their lids or caps.

        - Do not let the lower part of a sterile pipet touch anything except the insides of sterile containers.

        - Do not put container lids or caps on the table. Your instructor will show you how to hold them.

        - It is useful when working under aseptic conditions to think of the rest of the world as wet paint and the only areas that are not covered with wet paint are the surfaces and areas that have been sterilized by a deliberate process.

        - Eliminate or minimize talking. When talking, you are spewing out tiny droplets of saliva that are full of microorganisms.

      - Part 2 - Instructor Demonstration and Student Practice

        - Watch as your instructor demonstrates aseptic transfer by pipetting, including the do’s and don’ts.

        - Use your bottle of tap water to practice pipetting 5 mL into a test tube. Practice until you are comfortable with the technique. (You can use the same test tube over and over.)

        - Watch as your instructor demonstrates aseptic transfer by pouring, including the do’s and don’ts.

        - Use your bottle of tap water to practice pouring 10 mL into a 50 mL test tube or 50 mL Bottle. Practice until you are comfortable with the technique. (You can use the same test tube over and over.)

      - Part 3 - Aseptic Transfer by Pipetting

        - Before beginning work

- Wash your hands.

- Disinfect the work area of your bench with 70% isopropanol.

- Put on gloves and disinfect them with 70% isopropanol.

        - Use aseptic technique to pipet 5 ml of sterile medium into each of 6 tubes.

        - Label each tube and the bottle of remaining growth medium with your initials and the date.

        - Put tubes and bottle into the incubator (37°C) as directed by your instructor.

| Document No.: BioWork-7030 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 4 of 5 |
| --- | --- | --- | --- |
| Aseptic Transfers on an Open Bench |  |  |  |

      - Part 4 - Aseptic Transfer by Pouring

        - Start with a new bottle of sterile medium.

        - Use aseptic technique to pour 10 mL of sterile medium into each of 4 tubes (50 mL Conical Tubes or 50 mL sterile bottles).Be careful not to let liquid dribble down the outside neck of the bottle and then be transferred to a new sterile test tube.

Note: Wipe spills immediately with 70% IPA Wipes or equivalent

        - Label each tube and the bottle of remaining growth medium with your initials and the date.

        - Put tubes and bottle into the incubator (37°C) as directed by your instructor.

      - Evaluating Your Aseptic Technique

        - After the incubation period (24 – 72 hrs):

          - Retrieve your materials from the incubator.

          - Check each container for contamination. Look for turbidity or cloudiness in liquids. It is often useful to swirl the contents of the bottle or tube to detect very low levels of turbidity.

          - Note the results below. Enter a checkmark for each container that is free of contamination and an [X] for each in which you see contamination.

N/A

    - Figures and Diagrams

| Document No.: BioWork-7030 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 5 of 5 |
| --- | --- | --- | --- |
| Aseptic Transfers on an Open Bench |  |  |  |

    - Notes/Observations

| Aseptic transfer by pipetting |  |  |  |
| --- | --- | --- | --- |
| Tube 1 | ☐ | Tube 4 | ☐ |
| Tube 2 | ☐ | Tube 5 | ☐ |
| Tube 3 | ☐ | Tube 6 | ☐ |
| Stock Bottle | ☐ |  |  |

| Aseptic transfer by pouring |  |  |  |
| --- | --- | --- | --- |
| Tube 1 | ☐ | Tube 3 | ☐ |
| Tube 2 | ☐ | Tube 4 | ☐ |
| Stock Bottle | ☐ |  |  |

| ✓ | If you checked every box, congratulations! |
| --- | --- |
| X | If you saw contamination in any containers from an experiment, you may have the time to repeat that experiment. Do not feel bad. Aseptic technique takes practice. |

    - Attachments

N/A

| Document No.: BioWork-7040 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 1 of 7 |
| --- | --- | --- | --- |
| Sterilizing Liquids—Autoclaving and Filtration |  |  |  |

#### Purpose

To demonstrate two of the methods commonly used to sterilize liquid solutions. To practice and observe the results of two different methods used to sterilize liquids—autoclaving and filtration—and to learn autoclave operation and safety.

#### Scope

This exercise has three parts. A microbial growth medium will be prepared and split into three samples (fractions). Each sample will be checked and compared for contamination.

- One fraction will be autoclaved.

- One fraction will be filter sterilized.

- One fraction will be left untreated to demonstrate there was bioburden in the sample and that autoclaving or filtration did eliminate that viable bioburden.

Note: This is the note section.

#### Responsibilities

Technicians will operate equipment under the direction of an instructor. It is the responsibility of students undergoing training to comply with this procedure.

#### Definitions

Microbial growth medium:	a nutrient source that will support the growth of microbes

Sterile:	free from bacteria or other living microorganisms

#### References

- Specific operating manual for autoclave at each facility

#### Materials

| Material Description | Standard Quantity |  | Unit |
| --- | --- | --- | --- |
| Magnetic stirrer/hotplate | 1 |  | Each |
| Magnetic stir bar and retriever | 1 |  | Each |
| Top-loading balance | 1 |  | Each |
| Filtration unit (115 ml capacity, 0.2 µ pore size) | 1 |  | Each |
| Notes: |  |  |  |
| Notes: |  |  |  |
|  |  | BioWork Lab |  |
|  |  | BioWork Lab |  |

| Document No.: BioWork-7040 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 2 of 7 |
| --- | --- | --- | --- |
| Sterilizing Liquids—Autoclaving and Filtration |  |  |  |

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Trypticase soy broth (TSB) powder | 1 | Each |
| Weigh boat | 1 | Each |
| DI water in squirt bottle | 150 | mL |
| Spatula | 1 | Each |
| Autoclave tape | 2 | Pcs |
| Water-line aspirator or vacuum pump; (share with another pair of students) | 1 | Each |
| Tubing | 1-2 | Feet |
| 150-mL beaker | 1 | Each |
| 200-mL graduated cylinder | 1 | Each |
| 50 or 100-mL glass bottles with screw caps | 2 | Each |
| Heat-resistant gloves | 1 | Pair |
| Spray can of foaming alcohol or 70% isopropyl alcohol (IPA) solution | 1 | Each |
| Label tape | 1 | Each |
| Marking pen | 1 | Each |

- Safety

Through the entire activity, wear completely fastened lab coats, safety glasses, latex or nitrile gloves, and observe safety guidelines for operation of the autoclave.

Note: The autoclave will be hot! Don’t touch the chamber door or other surfaces while the autoclave is in use.

| Document No.: BioWork-7040 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 3 of 7 |
| --- | --- | --- | --- |
| Sterilizing Liquids—Autoclaving and Filtration |  |  |  |

#### Procedure

Note: Observe as your instructor demonstrates the following guidelines for safe operation of the autoclave.

  - General Instructions (Autoclave).

    - Put liquids in containers of glass or high-temperature plastic that can withstand heat. Do not fill containers completely full with liquids. Leave room in the container—about 1/5th of the volume—for liquids to expand.

    - Wrap small instruments or equipment in foil or paper so they remain sterile when removed from the autoclave.

    - Loosen screw caps and stoppers on glassware. Tightly closed containers can break or be crushed.

    - Stick pieces of autoclave indicator tape to items in the load. This tape develops distinctive markings when exposed to the right temperature for the minimum sterilization time. As an option, you can use a chemical indicator strip which will also show how long the items have been exposed to the sterilization temperature.

    - Do not pack materials in the chamber too tightly because the steam has to circulate freely and contact everything in the autoclave.

    - Follow required sterilization times. Large loads or volumes of liquids take a while to reach the right temperature for sterilization

    - Each autoclave will have specific operating instructions. Make sure you know what they are before starting the autoclave.

    - If the autoclave has an automatic cycle recorder, check to make sure that it is set up properly. If not, record temperature, pressure, or other parameters in a batch record as instructed.

    - Choose the appropriate exhaust cycle—fast for anything not containing liquids or slow for liquid loads.

      - Fast exhaust means that at the end of the sterilization time a valve is opened to vent steam from the chamber until the pressure in the chamber is equal to atmospheric pressure.

      - Slow exhaust means the steam supply shuts off and the chamber cools down slowly. As the chamber cools, the steam condenses and pressure drops slowly to atmospheric pressure or below. You must use slow exhaust when sterilizing liquids.

    - Never attempt to open the chamber door unless the pressure is at zero. (Many autoclaves have safety interlocks to prevent this.)

    - Before opening the chamber door, put on heat-resistant gloves. Stand to one side so that steam in the chamber will not hit you in the face.

    - When unloading liquids, handle the containers gently. If a liquid is swirled or agitated before it has cooled completely, it can suddenly boil up and out of the container and burn you.

| Document No.: BioWork-7040 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 4 of 7 |
| --- | --- | --- | --- |
| Sterilizing Liquids—Autoclaving and Filtration |  |  |  |

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Weigh out 3.6 g of soy broth powder in a clean weigh boat |  |  |
| 2. | Transfer the soy broth powder to a 150-mL beaker |  |  |
| 3. | Add about 100 mL of water to the beaker |  |  |
| 4. | Mix on a magnetic stirrer/hotplate until dissolved |  |  |
| 5. | Pour the mixture into a 200-mL graduated cylinder and add water to a final volume of 120 ml.; Swirl to mix |  |  |
| 6. | Divide the TSB into three 40-ml parts: - Pour 40 ml of broth into one bottle, cap it, and stick a short piece of autoclave tape over the cap. Write control, your initials, and date on a piece of label tape and put the tape on the bottle. - Pour another 40 mL of broth into a second bottle, loosely cap it, and stick a short piece of autoclave tape over the cap. Write AUTOCLAVED, your initials, and date on a piece of label tape and put the tape on the bottle. - Leave the remaining 40 ml of broth in the graduated cylinder for filtration. Note: See Flow Diagram |  |  |
| 7. | Follow the autoclave operation and safety guidelines in part 1 and the instructor’s directions to autoclave the sample marked autoclaved. |  |  |

Preparing and Sterilizing Samples for Incubation Prepare Broth (TSB)

| Document No.: BioWork-7040 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 5 of 7 |
| --- | --- | --- | --- |
| Sterilizing Liquids—Autoclaving and Filtration |  |  |  |

Sterile Filtration

| Step | Process Description | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- |
| 1. | Filter the broth in the graduated cylinder. Your instructor will demonstrate how to use the filtration unit. |  |  |
| 2. | Disinfect gloved hands with foaming alcohol or 70% IPA. |  |  |
| 3. | Turn on the pump or water line and then attach the tubing to the sidearm of the filtration unit. |  |  |
| 4. | Make sure that the cotton filter on the sidearm is in place and that the filtration unit is stable. Hold it upright if necessary. The cotton filter prevents contamination from air that rushes into the lower chamber when the vacuum is released. |  |  |
| 5. | Remove the lid from the upper chamber of the unit, pour your TSB into the chamber, and replace the lid. |  |  |
| 6. | When all the liquid has gone through the filter in the bottom of the chamber, hold the filtration unit firmly and detach the vacuum line from the sidearm. Be careful not to dislodge the cotton filter on the sidearm. |  |  |
| 7. | Turn off the pump or water line. |  |  |
| 8. | Label the filtration unit with your initials and date. |  |  |
| 9. | After autoclaving is complete, collect the three samples of TSB. |  |  |
| 10 | Place them in the incubator at 37°C as directed by your instructor |  |  |

  - After 24-hour Incubation

    - Remove your containers of soy broth from the incubator.

    - Hold the containers up to the light and gently swirl the liquid. A cloudy liquid means the liquid is contaminated (not sterile).

    - Record your observations below and discuss with the class.

| Document No.: BioWork-7040 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 6 of 7 |
| --- | --- | --- | --- |
| Sterilizing Liquids—Autoclaving and Filtration |  |  |  |

#### Figures and Diagrams

Preparation Flow Diagram

Autoclave Diagram

P

Steam to

Chamber

Steam

Chamber

Air

Door

Steam to Jacket

Condensate

T

### to waste

| Document No.: BioWork-7040 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 7 of 7 |
| --- | --- | --- | --- |
| Sterilizing Liquids—Autoclaving and Filtration |  |  |  |

#### Notes/Observations

Observations

| Sample | Observations |
| --- | --- |
| Control (not sterilized) |  |
| Autoclaved |  |
| Filtered |  |

Answer the questions that follow and discuss with the class:

| If you observe cloudiness in the autoclaved or filtered samples, think back over your work and suggest reasons why these samples might have become contaminated. |  |
| --- | --- |
| Even if the filtered sample is clear, do you think there could still be contamination in this sample? Why? |  |

N/A

#### Attachments

| Document No.: BioWork-7050 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 1 of 6 |
| --- | --- | --- | --- |
| Aseptic Transfers in an Laminar Airflow Hood (simulated) |  |  |  |

- Purpose

To learn proper aseptic technique for transferring sterile liquids or cultures from one container to another—one of the most common tasks you will perform. To perform these activities with the simulated restrictions of a Biosafety Cabinet (Laminar Airflow Hood)

#### Scope

In this activity, you will practice aseptic transfer techniques for pipetting and pouring sterile growth media in a simulated laminar flow hood. After the transferred media has incubated, you will evaluate your aseptic transfer technique by observing the media for contamination. The activity has six parts:

- Prep: Construction of a simulated Biosafety Cabinet

- Part 1: Precautions for Aseptic Transfer

- Part 2: Instructor Demonstration and Student Practice

- Part 3: Aseptic Transfer by Pipetting

- Part 4: Aseptic Transfer by Pouring

- Part 5: Evaluating Your Aseptic Technique

#### Responsibilities

Technicians will operate equipment and conduct the exercise under the direction of an instructor. It is the responsibility of students undergoing training to comply with this procedure.

#### Definitions

Aseptic:	Free from contamination with microorganisms. Asepsis or aseptic means free from pathogenic microorganisms.

Sterile:	Free from bacteria or other microorganisms Sterility can be defined as the freedom from the presence of viable microorganisms. Sterilization is the process of eliminating living microorganisms of all shapes and forms.

Turbidity: The quality of being cloudy, opaque, or thick with suspended mater.

#### References

- “Guideline on Sterile Drug Products Produced by Aseptic Processing,” Center for Drugs and Biologics and Office of Regulatory Affairs, Food and Drug Administration, Sept. 2004.

https://www.fda.gov/regulatory-information/search-fda-guidance-documents/sterile-drug- products- produced-aseptic-processing-current-good-manufacturing-practice

| Document No.: BioWork-7050 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 2 of 6 |
| --- | --- | --- | --- |
| Aseptic Transfers in an Laminar Airflow Hood (simulated) |  |  |  |

#### Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Ring stands (15-24 inches tall) and side clamps (or twist ties) | 2 | each |
| String or one 3 ft. bamboo plant stake | 3 | Ft |
| 50-mL bottle of tap water | 1 | Each |
| 50-mL bottles of sterile Tryptic Soy Broth growth medium | 2 | Each |
| Sterile 10-mL pipets | 3 | Each |
| Automatic pipet aid (share with one other student) | 1 | Each |
| Sterile 15-mL test tubes (For Pipetting) | 8 | Each |
| Sterile 50-mL Bottles, or 50 mL pre-sterilized (plastic) tubes, (For Pours) | 5 | Each |
| Test tube rack | 2 | Each |
| Marking pen | 1 | Each |
| Labeling tape | 1 | Roll |
| Latex or nitrile gloves | 1 | Pair |
| Spray can of foaming alcohol or spray bottle of 70% Isopropanol (share with one other student) | 1 | Each |

- Safety

Through the entire activity, you must wear safety glasses, completely fastened lab coats, and latex gloves (nitrile if you are allergic to latex).

#### Procedure

  - Prep – Construction of Simulated Biosafety Cabinet

| Document No.: BioWork-7050 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 3 of 6 |
| --- | --- | --- | --- |
| Aseptic Transfers in an Laminar Airflow Hood (simulated) |  |  |  |

Note: See diagram below, Section 9

    - Using tape divide a worktop into a 3 foot workstation. Team members should be able to work side by side.

    - Obtain (2) ring stands and place them at the front edge of the workspace. Space the rind stands about 3 feet apart. If clamps are available, place side clamp on each ring stand at approximately 12-15 inches off the surface of the bench.

    - Place the plant stake across the ring stands so it is balanced on the side arms.

    - If no side clamps are available, suspend a string across both ring stands at a height approximately 12-15 inches off the surface of the bench. Secure the ends of the string with twist ties.

    - Perform the activities below with the following restrictions that apply to working in the hood:

- Keep all work and equipment to the back of the simulated workspace

- The plant stake or string represents the sash of the Biosafety cabinet which must be pulled down to a certain level to maintain safety. The operator must work through the opening in the sash. In this case underneath the stake or string.

- Only the arms from the elbows to the hands may enter the workspace. The rest of the body must stay out of the work zone

  - Part 1 – Precautions for Aseptic Transfer (General Instruction)

    - Work quickly to minimize the time that containers are open.

    - Hold containers on a slant, not straight up and down, so that there is less chance of air particles falling into the containers.

    - To suction liquid into a pipet, always use a pipet aid—never your mouth.

    - Do not touch the mouths of containers or insides of their lids or caps.

    - Do not let the lower part of a sterile pipet touch anything except the insides of sterile containers.

    - Do not put container lids or caps on the table. Your instructor will show you how to hold them.

    - It is useful when working under aseptic conditions to think of the rest of the world as wet paint and the only areas that are not covered with wet paint are the surfaces and areas that have been sterilized by a deliberate process.

| Document No.: BioWork-7050 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 4 of 6 |
| --- | --- | --- | --- |
| Aseptic Transfers in an Laminar Airflow Hood (simulated) |  |  |  |

    - Eliminate or minimize talking. When talking, you are spewing out tiny droplets of saliva that are full of microorganisms.

  - Part 2 - Instructor Demonstration and Student Practice

    - Watch as your instructor demonstrates aseptic transfer by pipetting, including the do’s and don’ts.

    - Use your bottle of tap water to practice pipetting 5 mL into a test tube. Practice until you are comfortable with the technique. (You can use the same test tube over and over.)

    - Watch as your instructor demonstrates aseptic transfer by pouring, including the do’s and don’ts.

    - Use your bottle of tap water to practice pouring 10 mL into a 50 mL test tube or 50 mL Bottle. Practice until you are comfortable with the technique. (You can use the same test tube over and over.)

  - Part 3 - Aseptic Transfer by Pipetting

    - Before beginning work

- Wash your hands.

- Disinfect the work area of your bench with 70% isopropanol.

- Put on gloves and disinfect them with 70% isopropanol.

    - Use aseptic technique to pipet 5 ml of sterile medium into each of 6 tubes.

    - Label each tube and the bottle of remaining growth medium with your initials and the date.

    - Put tubes and bottle into the incubator (37°C) as directed by your instructor.

  - Part 4 - Aseptic Transfer by Pouring

    - Start with a new bottle of sterile medium.

    - Use aseptic technique to pour 10 mL of sterile medium into each of 4 tubes (50 mL Conical Tubes or 50 mL sterile bottles).Be careful not to let liquid dribble down the outside neck of the bottle and then be transferred to a new sterile test tube.

Note: Wipe spills immediately with 70% IPA Wipes or equivalent

    - Label each tube and the bottle of remaining growth medium with your initials and the date.

| Document No.: BioWork-7050 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 5 of 6 |
| --- | --- | --- | --- |
| Aseptic Transfers in an Laminar Airflow Hood (simulated) |  |  |  |

    - Put tubes and bottle into the incubator (37°C) as directed by your instructor.

  - Evaluating Your Aseptic Technique

    - After the incubation period (24 – 72 hrs.)

      - Retrieve your materials from the incubator.

      - Check each container for contamination. Look for turbidity or cloudiness in liquids. It is often useful to swirl the contents of the bottle or tube to detect very low levels of turbidity.

      - Note the results below. Enter a checkmark for each container that is free of contamination and an [X] for each in which you see contamination.

#### Figures and Diagrams

Simulated Biosafety Cabinet Work Space

3 feet

Simulated Sash

Simulated Bench Space

Ring Stands

| Document No.: BioWork-7050 | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 6 of 6 |
| --- | --- | --- | --- |
| Aseptic Transfers in an Laminar Airflow Hood (simulated) |  |  |  |

#### Notes/Observations

| Aseptic transfer by pipetting |  |  |  |
| --- | --- | --- | --- |
| Tube 1 | ☐ | Tube 4 | ☐ |
| Tube 2 | ☐ | Tube 5 | ☐ |
| Tube 3 | ☐ | Tube 6 | ☐ |
| Stock Bottle | ☐ |  |  |

| Aseptic transfer by pouring |  |  |  |
| --- | --- | --- | --- |
| Tube 1 | ☐ | Tube 3 | ☐ |
| Tube 2 | ☐ | Tube 4 | ☐ |
| Stock Bottle | ☐ |  |  |

| ✓ | If you checked every box, congratulations! |
| --- | --- |
| X | If you saw contamination in any containers from an experiment, you may have the time to repeat that experiment. Do not feel bad. Aseptic technique takes practice. |

- Attachments

N/A

| Document No.: BioWork-7060-BPR | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 1 of 13 |
| --- | --- | --- | --- |
| Detergent Experiment – Disinfecting Surfaces with Chemicals |  |  |  |

Laboratory / Process Description:

Your instructor has contaminated and taped off four squares (designated ‘A’ through ‘D’) on the workbench in the laboratory. Each square is contaminated with “harmless” microbes, but you should be cautious to not touch the squares and to wear proper PPE (lab coat, gloves, and safety glasses) throughout the lab. Each square A-D is contaminated with a genetically “crippled,” harmless strain of Escherichia coli. If the spore option is conducted, Squares (E-H) have been contaminated with Bacillus thuringiensis spores.

In this activity, you will see how disinfectants and detergents are used to eliminate microbes in a working environment and compare results based on the cleaning procedures performed. The following will be observed:

- Baseline – starting point for comparisons.

- Disinfectant only (70% IPA) – Contaminated surfaces wiped with a disinfectant only.

- Detergent only – Contaminated surfaces cleaned with detergent only.

- Detergent + Disinfectant – Contaminated surfaces cleaned with detergent and disinfectant.

Optional Spore Experiment:

- Baseline – starting point for comparisons.

- Disinfectant only (70% IPA) – Contaminated surfaces wiped with a disinfectant only.

- Hydrogen Peroxide only – Contaminated surfaces wiped with Hydrogen Peroxide only.

- Bleach (5%) only – Contaminated surfaces wiped with 5% Bleach only.

Other options for the experiment if available include Vesphene or LPh (Industry disinfectants);

DeconClean (Industry pH neutral detergent); and SporKlenz (Industry Sporicide).

This activity will be performed in two parts:

Part 1 – Performing the lab and incubation of samples.

Part 2 – Sample observations at your next class session.

Signature Log

| Name (print) | Signature | Initials | Date |
| --- | --- | --- | --- |

| Document No.: BioWork-7060-BPR | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 2 of 13 |
| --- | --- | --- | --- |
| Detergent Experiment – Disinfecting Surfaces with Chemicals |  |  |  |

Bill of Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| 2% Detergent Solution | 1 | Each |
| 70% Isopropyl Alcohol | 1 | Each |
| Hydrogen Peroxide (3% Drug Store bought) | 1 | Each |
| 5% Household Bleach Solution | 1 | Each |
| Lab Marker | 1 | Each |
| Lab Tape | 1 | Each |
| Low shedding wiper (optional Foam / sponge 1 inch paint brushes) | NS | Each |
| Sterile Swabs | 4-8 | Each |
| Timer | 1 | Each |
| Tryptic Soy Agar (TSA) Settling Plate | 4-8 | Each |
| Vials of Sterile Water | 4-8 | Each |
| For Instructor Set-Up |  |  |
| Fresh broth culture of non-pathogenic culture of E. coli | NS | Each |
| Diluted stock of Bacillus thuringiensis spores (Can be purchased through Home Depot: “Monterey B.t.”; for organic gardeners; 1 pint); must be diluted 1/1000 in sterile water.) | NS | Each |
| Industrial Disinfectants (If Available) | NS | Each |

#### *NS – Not specified

| Document No.: BioWork-7060-BPR | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 3 of 13 |
| --- | --- | --- | --- |
| Detergent Experiment – Disinfecting Surfaces with Chemicals |  |  |  |

Equipment List

| Equipment Description | Standard Quantity | Unit |
| --- | --- | --- |
| Incubator set at 32.5C ± 2.5C | 1 | Each |

Room Conditions (OPTIONAL)

| Room Conditions Check: Room |  |  | Performed By / Date | Verified By / Date |
| --- | --- | --- | --- | --- |
| Temp: °C | RH: % | ∆P: “WC |  |  |

| Document No.: BioWork-7060-BPR | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 4 of 13 |
| --- | --- | --- | --- |
| Detergent Experiment – Disinfecting Surfaces with Chemicals |  |  |  |

Procedure

Instructions for Taking a Swab Sample

| Step | Process Description |
| --- | --- |
| 1. | Dip a sterile swab in a vial of sterile water and expel water from the swab by pressing the swab against the inside of the vial. |
| 2. | Starting at the top left corner of a contaminated square, rub the swab across the square to the bottom right corner. Be sure to rotate the swab a ¼ turn as you collect your sample and not to touch the tape! |
| 3. | Open the TSA Settling plate and rub the swab evenly across the entire surface of the plate, rotating the swab frequently, then close the plate. Do not rub the swab on any portion of the plate more than once. |
| 4. | Dispose of the swab in a biohazard trash container. |

Instructions for Detergent Experiment – Part 1 (Escherichia coli)

| Step | Process Description |  |  |  | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- | --- | --- |
| 1. | Ensure that your workstation is set-up as illustrated. |  |  |  |  |  |
|  |  | A | B |  |  |  |
|  |  | C | D |  |  |  |
| 2. | Using a lab marker, label the base of each TSA Settling Plate (agar side up) with the square letter, cleaning step, initials and date. For example: - Baseline – Initials and date - Disinfectant Only (70% IPA) – Initials and date - Detergent Only – Initials and date - Detergent + Disinfectant – Initials and date |  |  |  |  |  |
| 3. | Using “Instructions for Taking a Swab Sample”, take a swab sample of Square A. |  |  |  |  |  |

| Document No.: BioWork-7060-BPR | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 5 of 13 |
| --- | --- | --- | --- |
| Detergent Experiment – Disinfecting Surfaces with Chemicals |  |  |  |

| Step | Process Description |  |  |  | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- | --- | --- |
| 4. | Spray Square B with 70% Isopropyl Alcohol and allow to dry. |  |  |  |  |  |
| 5. | Using “Instructions for Taking a Swab Sample”, take a swab sample of Square B. |  |  |  |  |  |
| 6. | Using a low shedding wiper, carefully wipe Square C with the 2% Detergent solution. Allow Square C to be saturated with the detergent solution for 2 minutes, then wipe dry. |  |  |  |  |  |
| 7. | Using “Instructions for Taking a Swab Sample”, take a swab sample of Square C. |  |  |  |  |  |
| 8. | Using a low shedding wiper, carefully wipe Square D with the 2% detergent solution. Allow Square D to be saturated with the detergent solution for 2 minutes, then wipe dry. |  |  |  |  |  |
| 9. | Spray Square D with 70% Isopropyl Alcohol and allow to dry. |  |  |  |  |  |
| 10. | Using “Instructions for Taking a Swab Sample”, take a swab sample of Square D. |  |  |  |  |  |
| 11. | Stack plates together and wrap once with tape. Incubate the samples for at least 36 hours at 32.5  2.5C. Record the following: |  |  |  |  |  |
|  |  | Incubation Information |  |  |  |  |
|  |  | Equipment ID |  |  |  |  |
|  |  | Temperature | C |  |  |  |
|  |  | Time In |  |  |  |  |
|  |  | Number of Samples |  |  |  |  |
| 12. | Remove the tape from the workbench and clean the area as instructed by your lab instructor. |  |  |  |  |  |

| Document No.: BioWork-7060-BPR | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 6 of 13 |
| --- | --- | --- | --- |
| Detergent Experiment – Disinfecting Surfaces with Chemicals |  |  |  |

Instructions for Detergent Experiment – Optional Part 1 (Bacillus thuringiensis spores)

| Step | Process Description |  |  |  | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- | --- | --- |
| 1. | Ensure that your workstation is set-up as illustrated. |  |  |  |  |  |
|  |  | E | F |  |  |  |
|  |  | G | H |  |  |  |
| 2. | Using a lab marker, label the base of each TSA Settling Plate (agar side up) with the square letter, cleaning step, initials and date. For example: - Baseline – Initials and date - Disinfectant Only (IPA) – Initials and date - Hydrogen Peroxide Only – Initials and date - Bleach Only – Initials and date |  |  |  |  |  |
| 3. | Using “Instructions for Taking a Swab Sample”, take a swab sample of Square E. (Baseline) |  |  |  |  |  |
| 4. | Lightly spritz Square F with 70% IPA. Using a low shedding wiper or foam brush, LIGHTLY wipe the square to spread the 70% IPA evenly. Start the timer set for 4 minutes. After the 4 minutes elapses, lightly spritz the square with 70% IPA. Again LIGHTLY wipe the square to spread the 70% IPA evenly. Start the timer set for 4 minutes. After the 4 minutes elapses take a swab of Square F. |  |  |  |  |  |
| 5. | Using “Instructions for Taking a Swab Sample”, take a swab sample of Square F. |  |  |  |  |  |
| 6. | Lightly spritz Square G with Hydrogen Peroxide. Using a low shedding wiper or foam brush, LIGHTLY wipe the square to spread the Hydrogen Peroxide evenly. Start the timer set for 4 minutes. After the 4 minutes elapses, lightly spritz the square with Hydrogen Peroxide again. Again LIGHTLY wipe the square to spread the Hydrogen Peroxide evenly. Start the timer set for 4 minutes. After the 4 minutes elapses take a swab of Square G. |  |  |  |  |  |

| Document No.: BioWork-7060-BPR |  |  | Revision No.: 02 |  | Effective Date: 29 Aug 2019 |  |  | Page 7 of 13 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Detergent Experiment – Disinfecting Surfaces with Chemicals |  |  |  |  |  |  |  |  |  |
| Step | Process Description |  |  |  |  |  | Performed By / Date: |  | Verified By / Date: |
| 7. | Using “Instructions for Taking a Swab Sample”, take a swab sample of Square G. |  |  |  |  |  |  |  |  |
| 8. | Lightly spritz Square H with 5% Bleach. Using a low shedding wiper or foam brush, LIGHTLY wipe the square to spread the 5% Bleach evenly. Start the timer set for 4 minutes. After the 4 minutes elapses, lightly spritz the square with 5% Bleach. Again LIGHTLY wipe the square to spread the 5% Bleach evenly. Start the timer set for 4 minutes. After the 4 minutes elapses take a swab of Square H. |  |  |  |  |  |  |  |  |
| 9. | Using “Instructions for Taking a Swab Sample”, take a swab sample of Square H. |  |  |  |  |  |  |  |  |
| 10. | Stack plates together and wrap once with tape. Incubate the samples for at least 36 hours at 32.5  2.5C. Record the following: |  |  |  |  |  |  |  |  |
|  |  | Incubation Information |  |  |  |  |  |  |  |
|  |  | Equipment ID |  |  |  |  |  |  |  |
|  |  | Temperature |  | C |  |  |  |  |  |
|  |  | Time In |  |  |  |  |  |  |  |
|  |  | Number of Samples |  |  |  |  |  |  |  |
| 11. | Remove the tape from the workbench and clean the area as instructed by your lab instructor. |  |  |  |  |  |  |  |  |

| Document No.: BioWork-7060-BPR | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 8 of 13 |
| --- | --- | --- | --- |
| Detergent Experiment – Disinfecting Surfaces with Chemicals |  |  |  |

Instructions for Detergent Experiment – Part 2 (Escherichia coli)

| Step | Process Description |  |  |  | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- | --- | --- |
| 1. | Remove your samples from the incubator after the required incubation times and record the following: |  |  |  |  |  |
|  |  | Incubation Information |  |  |  |  |
|  |  | Equipment ID |  |  |  |  |
|  |  | Temperature | C |  |  |  |
|  |  | Time Out |  |  |  |  |
|  |  | Number of Samples |  |  |  |  |
| 2. | Illustrate what you observe on Plate A: |  |  |  |  |  |
| 3. | Illustrate what you observe on Plate B: |  |  |  |  |  |

| Document No.: BioWork-7060-BPR | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 9 of 13 |
| --- | --- | --- | --- |
| Detergent Experiment – Disinfecting Surfaces with Chemicals |  |  |  |

| 4. | Illustrate what you observe on Plate C: |  |  |
| --- | --- | --- | --- |
| 5. | Illustrate what you observe on Plate D: |  |  |
| 6. | Discard plates in a biohazard bin and clean work area as instructed by your lab instructor. |  |  |

| Document No.: BioWork-7060-BPR | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 10 of 13 |
| --- | --- | --- | --- |
| Detergent Experiment – Disinfecting Surfaces with Chemicals |  |  |  |

Instructions for Detergent Experiment – Optional Part 2 (Bacillus thuringiensis spores)

| Step | Process Description |  |  |  | Performed By / Date: | Verified By / Date: |
| --- | --- | --- | --- | --- | --- | --- |
| 1. | Remove your samples from the incubator after the required incubation times and record the following: |  |  |  |  |  |
|  |  | Incubation Information |  |  |  |  |
|  |  | Equipment ID |  |  |  |  |
|  |  | Temperature | C |  |  |  |
|  |  | Time Out |  |  |  |  |
|  |  | Number of Samples |  |  |  |  |
| 2. | Illustrate what you observe on Plate E: |  |  |  |  |  |
| 3. | Illustrate what you observe on Plate F: |  |  |  |  |  |

| Document No.: BioWork-7060-BPR |  | Revision No.: 02 | Effective Date: 29 Aug 2019 |  | Page 11 of 13 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Detergent Experiment – Disinfecting Surfaces with Chemicals |  |  |  |  |  |  |  |
| 4. | Illustrate what you observe on Plate G: |  |  |  |  |  |  |
| 5. | Illustrate what you observe on Plate H: |  |  |  |  |  |  |
| 6. | Discard plates in a biohazard bin and clean work area as instructed by your lab instructor. |  |  |  |  |  |  |

| Document No.: BioWork-7060-BPR | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 12 of 13 |
| --- | --- | --- | --- |
| Detergent Experiment – Disinfecting Surfaces with Chemicals |  |  |  |

Equipment Printouts and/or Attachments:

Notes/Observations: Instructor’s Notes

E.Coli Stock Prep

- Pull 1 loop from plate of E.Coli, disperse into 50mL sterile TSB in a shaker flask.

- Shake overnight at 37°C “Night Culture”

- Using a sterile pipette, Take 5mL of the “Night Culture” and disperse into 50mL fresh sterile TSB. (Early on the morning of class)

- Shake 3-4 hours at 37°C

- Use (2) sterile swab to paint the test square (in both directions).

- Do 15 minutes before class to allow drying time

B. Thuringiensis Spore Prep

- Remove 10 mL Stock BT Spore solution (Monterey BT, Biological Insecticide)and add to 9mL sterile DI water or TSB for (10-1)dilution

- Remove 10 mL of the (10-1)Spore dilution and add to 9mL sterile DI or TSB for (10-2)dilution

- Remove 10 mL of the (10-2)Spore dilution and add to 9mL sterile DI or TSB for (10-3)dilution

- Use (2) sterile swab to paint the test square (in both directions) using the (10-3)dilution.

- Do 15 minutes before class to allow drying time

| Document No.: BioWork-7060-BPR | Revision No.: 02 | Effective Date: 29 Aug 2019 | Page 13 of 13 |
| --- | --- | --- | --- |
| Detergent Experiment – Disinfecting Surfaces with Chemicals |  |  |  |

Discussion Questions – Part 1

- What do you think is the purpose of the baseline plate?

- Describe what you expect to see on each plate.

- Which plate do you expect to see no growth? Why?

Discussion Questions – Part 2

- Based on your response from the Discussion Questions – Part 1, did you see growth as expected? Why or why not?

- If not, what do you think could have went wrong? How could you improve?