# Purpose

To characterize the quantity and quality of nucleic acid samples using a Nanodrop spectrophotometer.

# Scope

Determining the amount of DNA or RNA from extractions and assessing sample quality is recommended before PCR and other library preparation procedures.

# Responsibilities

Trained personnel under the supervision of authorized staff/faculty.

# Definitions

- Nanodrop spectrophotometer: A device used for measuring the concentration and purity of nucleic acids (DNA, RNA) by analyzing the absorbance of ultraviolet (UV) light.

- PCR (Polymerase Chain Reaction): A technique used in molecular biology to amplify a single copy or a few copies of a piece of DNA, generating thousands to millions of copies of a particular DNA sequence.

- Library preparation: A process in next-generation sequencing where nucleic acids (DNA or RNA) are prepared for sequencing. This involves fragmenting the DNA/RNA, adding adapters, and amplifying the fragments.

- DNA (Deoxyribonucleic Acid): A molecule that carries the genetic instructions used in the growth, development, functioning, and reproduction of all known living organisms and many viruses.

- RNA (Ribonucleic Acid): A molecule essential for various biological roles in coding, decoding, regulation, and expression of genes.

- Sterile Deionized Water (SDW): Water that has been purified to remove ions and sterilized to remove any microorganisms.

- Kim wipes: Laboratory wipes used for cleaning optical equipment and other sensitive instruments.

- Microcentrifuge tube: Small tubes used in laboratories for centrifugation, storing, and handling small quantities of liquids.

- Biohazard bag: A container used to collect and dispose of materials that are potentially infectious or hazardous to human health.

- Sample quality: The assessment of the purity and integrity of a nucleic acid sample, important for ensuring reliable and accurate downstream applications such as PCR and sequencing.

# References

N/A

# Materials

(Any materials here)

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| Adjustable micro pipette and tips – (0.2 μl – 2.0 μl) | N/A | N/A |
| Thermoscientific Nanodrop Lite | 1 | N/A |
| Sterile Deionized Water (SDW) | N/A | N/A |
| Kim wipes | N/A | N/A |
| DNA samples | N/A | N/A |

# Safety

Wear closed toe shoes, safety glasses, and disposable gloves when working. Lab jacket recommended to avoid contamination.

# Procedure

Navigation

- Navigate through the menu using the four black circular buttons location at the bottom of the screen. The buttons option choices are located on the screen at the bottom, aligned with the buttons.

- While navigating the home screen, you will see multiple options of what type of sample you are measuring such as DNA (ds or ss), mRNA, or proteins.

- Select the desired type of sample you are measuring.

Blanking

- Ensure you do this step before and after measuring your sample to avoid contamination.

- Gently lift the lid all the way up.

- Grab your pipette (0.2 μl – 2.0 μl) and set the volume to 1 μl.

- Draw 1 μl of water from the microcentrifuge tube.

- Dispense the sample on the small silver plate.

- Gentle close the lid, press blank, and discard the pipette tip in a biohazard bag.

- Once successfully finished, open the lid, grab a Kim wipe, and wipe the silver plate thoroughly.

Measuring samples

- Attach another tip to your pipette and draw 1 μl of your sample.

- Gently dispense your sample on the silver plate and close the lid.

- Eject the tip in a biohazard bag, close the lid, and select the “Measure” Option on the screen.

- Once it finishes measuring, write down your findings on the sheet of paper located near the machine and wipe the silver plate with a Kim Wipe.

- Repeat steps 1-4 with next sample or clean with drop of SDW on Kim Wipe to clean.

# Figures and Diagrams

N/A

# Notes/Observations

# Attachments

N/A