# Purpose

Using the “MagMax” DNA Multi-Sample Ultra 2.0 kit for rapid purification and extraction of high-quality DNA with the buccal swab technique using the KingFisher Flex.

# Scope

To extract required DNA from buccal swabs while operating KingFisher Flex.

# Responsibilities

Those under the supervision of trained personnel.

# Definitions

- Buccal swab: The painless removal of a sample of cells from the inside of your mouth (cheek) for study.

- Sequencing: Determining the order of the four chemical building blocks - called "bases" - that make up the DNA molecule.

- Genotyping: The experimental procedure that identifies the differences in DNA sequence among individuals or populations.

- KingFisher Flex: An automated magnetic particle processor used for DNA extraction and purification.

- Micropipette Sets and Tips (P2, P20, P200, P1000): Precision instruments used to measure and transfer small volumes of liquids, specified by their maximum capacity in microliters.

- Plate Shaker: A device used to shake plates containing samples to ensure thorough mixing.

- Vortex: A device used to mix samples quickly by creating a vortex in the liquid.

- Incubator: A device used to maintain a constant temperature for the incubation of samples.

- Reservoir Trays: Containers used to hold reagents temporarily during pipetting.

- Conical Tubes (15 mL and 50 mL): Tubes used for mixing and preparing solutions, typically with a conical bottom to facilitate the collection of precipitates.

- Deepwell Plates: Plates with deep wells used for sample processing in high-throughput applications.

- Standard Plates: Plates with standard well depths used for various laboratory applications.

- Tip Comb: A comb-like structure used in conjunction with the KingFisher Flex to handle samples during magnetic particle processing.

- Ethanol (96-100% Molecular Lab Grade): Alcohol used for washing and purifying DNA.

- Enhancer Solution: A solution used to enhance the efficiency of DNA extraction.

- Proteinase K: An enzyme used to digest proteins and aid in DNA extraction.

- Binding Solution: A solution used to facilitate the binding of DNA to magnetic particles.

- DNA Binding Beads: Magnetic beads used to bind DNA during the extraction process.

- Wash I Solution: A solution used to wash the DNA-bound beads to remove impurities.

- Elution Solution: A solution used to release the purified DNA from the magnetic beads.

- Phosphate Buffer Solution (PBS) 1x at pH 7.4: A buffer solution used to maintain the pH and osmolarity of the samples.

# References

MagMAX DNA Multi-Sample Ultra 2.0 Kit Protocol (Pub. No. MAN0017205 Rev. D.0)

# Materials

| Material Description | Standard Quantity | Unit |
| --- | --- | --- |
| KingFisher Flex | 1 | N/A |
| Micropipette sets and tips: P2, P20, P200, P1000, multichannel pipettors | N/A | N/A |
| Plate shaker | 1 | N/A |
| Vortex | 1 | N/A |
| Incubator | 1 | N/A |
| Reservoir trays | N/A | N/A |
| 15 mL and 50 mL conical tubes | N/A | N/A |
| Deepwell plates | N/A | N/A |
| Standard plates | N/A | N/A |
| Tip comb compatible with KingFisher Flex | N/A | N/A |
| Ethanol (96-100%, molecular lab grade) | N/A | N/A |
| Enhancer Solution | N/A | N/A |
| Proteinase K | N/A | N/A |
| Binding Solution | N/A | N/A |
| DNA Binding Beads | N/A | N/A |
| Wash I Solution | N/A | N/A |
| Elution Solution | N/A | N/A |
| Phosphate Buffer Solution (PBS), 1x at pH 7.4 | N/A | N/A |
| Swabs with foam tip | N/A | N/A |

# Safety

- PPE (including gloves, safety glasses, and lab jacket) must be worn when following SOP. Gloves should be disinfected or changed frequently when handling buccal swabs.

- Dispose of lysis buffer waste in special waste container.

# Procedure

Before beginning

### Make sure the following program is loaded in the KingFisher: MagMAX_Ultra2_Direct_FLEX

### For first time use of kit, prepare Wash II Solution by making 80% ethanol from 100% ethanol and nuclease-free water. Initial the container (e.g. 15 or 50 mL conical tube) and timestamp the date of mixing.

### Before each use of kit, vortex DNA Binding Beads well.

### Check all the solutions in the kit for salt precipitation. Dissolve any precipitates by warming the solution at 37°C, then equilibrate to room temperature.

### Set heat block, water bath, or thermomixer to 65°C.

### Ensure that the instrument is set up for processing with the deep-well magnetic head and has the standard 96-well heating block installed.

Sample collection and storage

### Have test subjects rinse mouths with water and swallow prior to swabbing.

### Remove swab from packaging and thoroughly swab both cheeks of test subject for 30 seconds each.

### Put back in pouch and store in freezer up to three weeks or proceed directly to breaking top of each swab sample into a well of sample plate.

### Label and date plate and include sample information (e.g. sample well ID, date of extraction, operator) in batch record / lab notebook.

Prepare Proteinase K mix and digest samples

### Prepare sufficient amount of Proteinase K Mix by first determining the number of

### samples you want to extract (n) and planning for 20% extra (n + 0.2n).

### In a 50 mL conical tube (label as Proteinase K Mix), add the following amount of

### Enhancer Solution: 40 µL * (n + 0.2n).

### In the same conical tube, add the following amount of 1x PBS: 400 µL * (n + 0.2).

### In the same conical tube, add the following amount of Proteinase K:

### 40 µL * (n + 0.2).

### Invert the closed tube to mix and pour in reservoir tray.

### Using multichannel or adjustable micropipette, add 480 µL of Proteinase K Mix into each well with a swab (be careful not to touch swabs with pipette tips).

### Seal the plate with clear adhesive film.

### Use plate shaker to shake the sealed plate at 900 rpm for 5 minutes.

### Place shaken plate in the incubator at 65° Celsius for 20 minutes and proceed to preparation of plated reagents while sample plate incubates.

Preparation of plated reagents

### Setup deep well plate of Wash I solution by dumping reagent in clean reservoir tray, pipetting with a multichannel 1000 µL of reagent into each well being used, and placing plate into position 2 of KingFisher Flex.

### Setup deep well plate of Wash II solution by dumping reagent in clean reservoir tray, pipetting with a multichannel 1000 µL of reagent into each well being used, and placing plate into position 3 of KingFisher Flex.

### Setup another deep well plate of Wash II solution by dumping reagent in clean reservoir tray, pipetting with a multichannel 500 µL of reagent into each well being used, and placing plate into position 4 of KingFisher Flex.

### Setup standard plate of Elution solution by dumping reagent in clean reservoir tray, pipetting with a multichannel 50 µL of reagent into each well being used, and placing plate into position 5 of KingFisher Flex.

### Place tip comb on standard plate in position 6 of KingFisher Flex instrument.

### In a 15 or 50 mL conical tube label (label as Binding Bead Mix), add the following amount of Binding Solution: 400 µL * (n + 0.2n).

### In the same conical tube, add the following amount of pre-vortexed DNA Binding Beads: 40 µL * (n + 0.2).

Prepare final sample plate

### After sample plate plus Proteinase K mixture has finished digesting in the incubator, transfer the sealed plate to the plate shaker for shaking at 900 rpm for 2 minutes.

### Transfer the lysates into the corresponding wells of new plate (label as sample plate and date) by setting multichannel to 480 µL and making sure at least 420 µL of lysates have been transferred. Change tips for each transfer of samples.

### For each well in the new sample plate containing transferred lysates, add 440 µL of DNA Binding Bead Mix by pouring mix in reservoir tray and using multichannel micropipette (remix binding bead mix frequently by slowly pipetting up and down before adding to samples; avoid creating air bubbles during mixing and aliquoting).

### Place new sample plate in position 1 of KingFisher Flex instrument (make sure A1 of all plates are correctly oriented with the placeholders of the KingFisher Flex instrument).

Run KingFisher Flex program

### Turn on KingFisher Flex with switch on the left side of machine.

### Use arrow buttons on panel to navigate to program

### “MagMAX_Ultra2_Direct_FLEX” program (go to middle tab on panel and scroll down) and hit ‘START’ button.

### Follow prompts of instrument panel to confirm proper placement of each reagent plate, tip comb, and sample plates, hitting ‘START’ between each prompt and once more to start program.

### At the end of the run, remove elution plate and transfer eluate to new plate or tubes (noting sample identity directly on tubes or on batch record matching well number of sample plate with the identity record of sample). Store eluted samples in freezer or proceed to assays of concentration and quality (refer to SOP VG005).

# Figures and Diagrams

N/A

# Notes/Observations

# Attachments

See DNA Extraction Batch Record and Accession-Sampling Sheet Template